SPD TECHNICIAN

TRAINING GUIDE

────────────────────────────────────────

MODULE 1

FOUNDATIONS

Why This Work Matters

IN THIS MODULE • The Mission: Why Sterile Processing Exists • The Science: Microbiology You Need to Know • The System: How Infections Spread • The Framework: Spaulding Classification • The Standards: Regulations That Govern Our Work • The Method: Sterile By Design™

SIPS HEALTHCARE SOLUTIONS

Sterile By Design™ Curriculum

📚 LEARNING OBJECTIVES

After completing this module, you will be able to: 1. Explain why sterile processing is critical to patient safety 2. Describe the basic types of microorganisms and their relevance to SPD 3. Explain the chain of infection and how SPD breaks it 4. Apply the Spaulding Classification system to categorize medical devices 5. Identify the key regulatory bodies and standards that govern sterile processing 6. Describe the four pillars of Sterile By Design™ and their application 7. Demonstrate the professional mindset required for this work

# 1.1 The Mission: Why Sterile Processing Exists

❤️ THE PATIENT CONNECTION Before we discuss procedures, equipment, or techniques, you need to understand one thing: Everything you do in this department directly affects whether a patient lives or dies, heals or suffers. This is not an exaggeration. This is the reality of your job.

## The Purpose of Sterile Processing

The Sterile Processing Department exists for one reason: to ensure that every reusable medical device used on a patient is safe — free from contamination that could cause infection.

Every day in hospitals across the country, thousands of surgical procedures take place. Each one requires instruments that have been used on other patients before. Your job is to take those contaminated instruments and make them safe to use again.

### What Happens If We Fail

When sterile processing fails, patients get infections. These are called Healthcare-Associated Infections (HAIs) or Surgical Site Infections (SSIs).

| THE REALITY OF SURGICAL SITE INFECTIONS |  |
| --- | --- |
| Patients affected annually (US) | ~300,000 surgical site infections per year |
| Additional hospital days | 7-10 extra days per infection |
| Cost per infection | $20,000 - $90,000 additional cost |
| Mortality | 2-11% of SSI patients die from the infection |
| Preventability | Up to 60% of SSIs are preventable |

These are not just statistics. Each number represents a real person — someone's parent, child, spouse, or friend — who suffered because something went wrong.

🎯 Your Role in Prevention You are the last line of defense before an instrument touches a patient. The surgeon trusts that the instruments are sterile. The patient trusts the surgeon. That chain of trust depends entirely on you doing your job correctly.

## The Invisible Heroes

Patients will never know your name. They will never see your face. When they wake up from surgery, they won't think about the person who cleaned, inspected, and sterilized the instruments used on them.

But you will know. Every tray you process correctly is a patient protected. Every error you catch is an infection prevented. Every shortcut you refuse to take is a life you may have saved.

### The Professional Mindset

This job requires a specific mindset:

- Attention to detail: Small things matter. A speck of tissue left on an instrument can harbor bacteria.

- Consistency: You must do things the same way every time, not just when someone is watching.

- Integrity: If you make a mistake, you report it. Covering up errors puts patients at risk.

- Humility: No matter how experienced you become, you follow the procedures.

- Pride: Take pride in your work. Your name on a tray means something.

# 1.2 The Science: Microbiology You Need to Know

You don't need to be a microbiologist to work in SPD. But you do need to understand the basics of what you're fighting against.

## What Are Microorganisms?

Microorganisms (also called microbes or "germs") are living things too small to see without a microscope. Some are harmless. Some are helpful. And some cause disease — these are called pathogens.

### Types of Microorganisms

| TYPE | WHAT IT IS | EXAMPLES | SPD RELEVANCE |
| --- | --- | --- | --- |
| BACTERIA | Single-celled organisms; some cause infection | Staph, Strep, E. coli, MRSA | Killed by proper cleaning and sterilization |
| VIRUSES | Tiny particles that need a host cell to reproduce | Hepatitis B & C, HIV, COVID-19 | Inactivated by cleaning and sterilization |
| FUNGI | Organisms including yeasts and molds | Candida, Aspergillus | Killed by standard sterilization |
| BACTERIAL SPORES | Dormant, highly resistant form of certain bacteria | Clostridium difficile, Bacillus | VERY hard to kill — requires sterilization |
| PRIONS | Abnormal proteins that cause brain disease | CJD (Creutzfeldt-Jakob disease) | Cannot be killed by standard sterilization — special handling required |

## Why This Matters to You

📌 The Hierarchy of Resistance Not all microorganisms are equally easy to kill. Understanding this hierarchy explains WHY we have different levels of processing: EASIEST TO KILL → → → → → → → → → HARDEST TO KILL Bacteria (vegetative) → Viruses → Fungi → Mycobacteria → Bacterial Spores → Prions Sterilization kills everything up to and including bacterial spores. Prions require special protocols beyond standard sterilization.

### Bacterial Spores: Your Primary Target

Bacterial spores deserve special attention because they are your benchmark. If your sterilization process can kill bacterial spores, it can kill everything else (except prions).

Spores are like survival pods for bacteria. When conditions become hostile, certain bacteria form spores — a dormant state with a tough outer shell that can survive:

- Extreme heat

- Extreme cold

- Drying

- Radiation

- Many chemicals

This is why we use Biological Indicators containing bacterial spores to test our sterilizers. If the sterilizer kills the spores, we know it's working.

## Bioburden: What We're Removing

Bioburden refers to the number and type of microorganisms on an item before processing. When instruments come from surgery, they have high bioburden — blood, tissue, body fluids, and the microorganisms that live in them.

### Why Cleaning Comes First

Sterilization cannot work on dirty instruments. Here's why:

- Organic matter shields microorganisms: Blood and tissue can protect bacteria from the sterilization process.

- Biofilm forms: Bacteria can form protective colonies that are extremely resistant to killing.

- Steam can't penetrate: If debris is present, steam cannot reach all surfaces.

⚠️ The Golden Rule You cannot sterilize a dirty instrument. Cleaning must ALWAYS come before sterilization. No exceptions.

# 1.3 The System: How Infections Spread

Understanding how infections spread helps you understand why each step in sterile processing matters.

## The Chain of Infection

For an infection to occur, six links must be present. This is called the Chain of Infection. Break any link, and you prevent the infection.

| LINK | WHAT IT MEANS | EXAMPLE |
| --- | --- | --- |
| 1. INFECTIOUS AGENT | The pathogen that causes disease | Staphylococcus aureus bacteria |
| 2. RESERVOIR | Where the pathogen lives and multiplies | Patient's wound, contaminated instrument |
| 3. PORTAL OF EXIT | How the pathogen leaves the reservoir | Blood on surgical instruments |
| 4. MODE OF TRANSMISSION | How the pathogen travels to a new host | Contaminated instrument used on next patient |
| 5. PORTAL OF ENTRY | How the pathogen enters the new host | Surgical incision |
| 6. SUSCEPTIBLE HOST | A person who can become infected | Surgical patient with open wound |

## How SPD Breaks the Chain

Your job is to break this chain — specifically at Link 4 (Mode of Transmission). You ensure that contaminated instruments do not become the vehicle that transmits infection from one patient to another.

| SPD ACTIVITY | HOW IT BREAKS THE CHAIN |
| --- | --- |
| Cleaning/Decontamination | Removes infectious agents and the organic matter they hide in |
| Inspection | Identifies any remaining contamination before sterilization |
| Sterilization | Kills all remaining microorganisms, including spores |
| Sterile Storage | Maintains sterility until the item is used |
| Proper Transport | Prevents recontamination during delivery |
| 🎯 You Are the Chain Breaker Think of yourself as a chain breaker. Your job is to ensure that no instrument becomes the link that transmits infection from one patient to another. Every step you take correctly is another break in that chain. |

# 1.4 The Framework: Spaulding Classification

Not all medical devices require the same level of processing. In 1968, Dr. Earle Spaulding developed a classification system that we still use today. Understanding this system tells you WHY we process different items differently.

## The Three Categories

| CATEGORY | CONTACT TYPE | INFECTION RISK | REQUIRED PROCESSING |
| --- | --- | --- | --- |
| CRITICAL | Enters sterile tissue or vascular system | HIGH | STERILIZATION |
| SEMI-CRITICAL | Contacts mucous membranes or non-intact skin | MODERATE | High-Level Disinfection (minimum) or Sterilization |
| NON-CRITICAL | Contacts intact skin only | LOW | Low-Level Disinfection |

## Critical Items

Critical items pose the highest risk of infection because they enter sterile body areas. These items MUST be sterile at the time of use.

### Examples of Critical Items

- Surgical instruments (scalpels, forceps, scissors, retractors)

- Implants (joint replacements, pacemakers, screws, plates)

- Cardiac catheters

- Needles

- Any item that enters the bloodstream or sterile tissue

⚠️ Critical = Sterile If an item is classified as CRITICAL, it must be STERILIZED. No exceptions. High-level disinfection is NOT acceptable for critical items.

## Semi-Critical Items

Semi-critical items contact mucous membranes or non-intact skin. Mucous membranes (found in the mouth, nose, throat, GI tract, etc.) are more resistant to infection than sterile tissue, but they can still be infected.

### Examples of Semi-Critical Items

- Flexible endoscopes (colonoscopes, bronchoscopes)

- Laryngoscope blades

- Respiratory therapy equipment

- Vaginal speculums

- Some probes and sensors

Semi-critical items require, at minimum, high-level disinfection (HLD). However, sterilization is preferred when the item can tolerate it.

## Non-Critical Items

Non-critical items contact only intact skin, which is an excellent barrier against infection. These items pose the lowest risk.

### Examples of Non-Critical Items

- Blood pressure cuffs

- Stethoscopes

- Bed rails

- Furniture

- Floors

Non-critical items require low-level disinfection — essentially, cleaning with a hospital-grade disinfectant.

## Applying the Spaulding System

💡 How to Classify Ask yourself: "What is the most invasive contact this item will have?" • Will it enter sterile tissue or the bloodstream? → CRITICAL → Sterilize • Will it touch mucous membranes or non-intact skin? → SEMI-CRITICAL → HLD or Sterilize • Will it touch only intact skin? → NON-CRITICAL → Low-level disinfection When in doubt, go with the higher level of processing.

# 1.5 The Standards: Regulations That Govern Our Work

Sterile processing doesn't operate in a vacuum. Multiple organizations set standards that govern what we do and how we do it. Understanding these helps you understand why we have so many rules.

## Key Regulatory and Standards Organizations

| ORGANIZATION | WHAT THEY DO | HOW IT AFFECTS YOU |
| --- | --- | --- |
| AAMI | Association for the Advancement of Medical Instrumentation — writes technical standards | Many facility policies are based on AAMI standards (ST79, ST91, etc.) |
| AORN | Association of periOperative Registered Nurses — sets perioperative guidelines | Provides guidelines for instrument processing in surgical settings |
| The Joint Commission (TJC) | Accredits healthcare organizations | Surveys hospitals and can cite deficiencies in SPD practices |
| CMS | Centers for Medicare & Medicaid Services — federal oversight | Sets Conditions of Participation; ties reimbursement to compliance |
| FDA | Food and Drug Administration — regulates medical devices | Requires following manufacturer IFUs; issues recalls |
| OSHA | Occupational Safety and Health Administration | Sets workplace safety rules including bloodborne pathogen standards |
| CDC | Centers for Disease Control and Prevention | Publishes infection control guidelines |
| State Health Departments | State-level regulation and inspection | May have additional requirements beyond federal standards |

## Key Standards You Should Know

### AAMI ST79

This is the primary standard for steam sterilization in healthcare. It covers decontamination, preparation, packaging, sterilization, and storage. Many of your facility's procedures are based on this standard.

### Manufacturer IFUs

IFU stands for Instructions for Use. These are the manufacturer's written instructions for how to process their specific device. Following IFUs is not optional — it's required by law.

⚠️ IFUs Take Precedence If a manufacturer's IFU is more stringent than your general policy, you follow the IFU. If you cannot follow the IFU (don't have the right equipment, for example), you cannot use that device until you can comply. Always check the IFU when processing an unfamiliar item.

## Why This Matters

When surveyors from The Joint Commission or your state health department visit, they will look at your SPD. They will check:

- Are you following IFUs?

- Are your sterilizers being tested correctly?

- Is documentation complete?

- Are staff trained and competent?

- Are policies up to date?

Deficiencies in SPD can result in citations, required corrective action plans, and in serious cases, can threaten the hospital's ability to operate.

Your compliance with procedures isn't just about patient safety (though that's the most important reason) — it's also about keeping the facility in good standing with regulators.

# 1.6 The Method: Sterile By Design™

🎯 STERILE BY DESIGN™ At SIPS, we don't leave sterility to chance. We design our processes so that doing things correctly is the natural outcome. Sterile By Design™ is not just a slogan — it's the operational philosophy that guides everything we do.

## The Four Pillars of Sterile By Design™

SBD rests on four pillars. Each pillar supports the others. Remove one, and the structure weakens.

### PILLAR 1: Standardization

STANDARDIZATION Same process, same way, every time, regardless of who is working.

Why it matters:

- Reduces variability — variability introduces error

- Ensures all staff produce the same quality outcome

- Makes training easier and more effective

- Allows problems to be identified and fixed

What it means for you:

- Follow procedures exactly as written

- Don't develop your own "better" way

- If you see a better way, propose it through proper channels — don't just do it

### PILLAR 2: Verification

VERIFICATION Systematic checking of work at defined points — never assume.

Why it matters:

- Catches errors before they reach the patient

- Creates accountability

- Builds quality into the process, not just at the end

What it means for you:

- Check your own work

- Use verification checkpoints built into procedures

- Never assume something was done — verify it

- When you verify, you're accountable for what you verified

### PILLAR 3: Documentation

DOCUMENTATION Complete, accurate, real-time records of all activities. If it isn't documented, it didn't happen.

Why it matters:

- Creates a traceable record for every item

- Supports investigation when problems occur

- Demonstrates compliance to regulators

- Protects patients, staff, and the facility

What it means for you:

- Document in real time — not at the end of your shift

- Document completely — no blank fields

- Document accurately — what actually happened

- Your signature/initials mean YOU verified it

### PILLAR 4: Communication

COMMUNICATION Standardized language to eliminate ambiguity. Say what you mean, mean what you say.

Why it matters:

- Prevents misunderstandings that lead to errors

- Ensures information transfers accurately between people

- Creates clear escalation pathways

What it means for you:

- Use the SBD standard phrases you'll learn

- Don't assume people know what you mean — be explicit

- Speak up when something isn't clear

- Report issues using the proper escalation language

# 1.7 The Commitment: Your Professional Responsibility

You are entering a profession. Not just a job — a profession with standards, ethics, and responsibilities.

## Professional Certification

Most SPD professionals pursue certification through one of two organizations:

| ORGANIZATION | CERTIFICATION | REQUIREMENTS |
| --- | --- | --- |
| HSPA | CRCST (Certified Registered Central Service Technician) | 400 hours work experience + exam |
| CBSPD | CSPDT (Certified Sterile Processing and Distribution Technician) | Experience requirements + exam |

Certification demonstrates that you have met a national standard of knowledge. Many employers require certification, and some states are moving toward making it mandatory.

## Professional Ethics

As a sterile processing professional, you have ethical obligations:

### To Patients

- Process every instrument as if it will be used on your own family

- Never take shortcuts that compromise patient safety

- Report errors that could affect patient care

### To Your Employer

- Follow policies and procedures

- Be honest in your documentation

- Report concerns through appropriate channels

### To Your Profession

- Maintain your competence through ongoing education

- Support your colleagues

- Represent the profession with integrity

### To Yourself

- Take pride in your work

- Set boundaries that protect your wellbeing

- Grow and develop professionally

# Knowledge Check: Module 1

Answer these questions to verify your understanding of this module. Discuss any questions with your preceptor.

### Section 1.1 — The Mission

- What is the primary purpose of the Sterile Processing Department?

_____________________________________________________________________________

_____________________________________________________________________________

- What are Healthcare-Associated Infections (HAIs) and why do they matter?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 1.2 — The Science

- List the five main types of microorganisms discussed in this module:

_____________________________________________________________________________

_____________________________________________________________________________

- Why are bacterial spores particularly important in sterile processing?

_____________________________________________________________________________

_____________________________________________________________________________

- Why must cleaning always come before sterilization?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 1.3 — How Infections Spread

- List the six links in the Chain of Infection:

_____________________________________________________________________________

_____________________________________________________________________________

- Which link in the chain does SPD primarily break? How?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 1.4 — Spaulding Classification

- Complete the table:

| CATEGORY | CONTACT TYPE | MINIMUM PROCESSING |
| --- | --- | --- |
| Critical |  |  |
| Semi-Critical |  |  |
| Non-Critical |  |  |

- Give one example of each category:

Critical: _________________________________________________________________

Semi-Critical: ____________________________________________________________

Non-Critical: _____________________________________________________________

### Section 1.5 — Regulations

- What does IFU stand for, and why is following IFUs mandatory?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 1.6 — Sterile By Design™

- List the four pillars of Sterile By Design™:

1. _________________________________________________________________________

2. _________________________________________________________________________

3. _________________________________________________________________________

4. _________________________________________________________________________

- What does "If it isn't documented, it didn't happen" mean in practice?

_____________________________________________________________________________

_____________________________________________________________________________

# Skills Validation: Module 1

This module focuses on foundational knowledge. Your preceptor will verify your understanding through discussion.

| VALIDATION ITEM | MET | NOT MET | DATE | INITIALS |
| --- | --- | --- | --- | --- |
| Learner can explain why sterile processing matters to patient safety | ☐ | ☐ |  |  |
| Learner can describe basic microorganism types and their resistance levels | ☐ | ☐ |  |  |
| Learner can explain the chain of infection and SPD's role in breaking it | ☐ | ☐ |  |  |
| Learner can correctly classify items using Spaulding system | ☐ | ☐ |  |  |
| Learner can explain why IFUs must be followed | ☐ | ☐ |  |  |
| Learner can state the four pillars of Sterile By Design™ | ☐ | ☐ |  |  |
| Learner completed Knowledge Check with satisfactory answers | ☐ | ☐ |  |  |

### Module Completion

| Learner Name: | ________________________________________ |
| --- | --- |
| Learner Signature: | ________________________________________ |
| Date: | ________________________________________ |
| Preceptor Name: | ________________________________________ |
| Preceptor Signature: | ________________________________________ |
| Date: | ________________________________________ |

# Module 1 Summary

✓ Key Takeaways ✓ Sterile processing exists to protect patients from infection ✓ Microorganisms have varying resistance — spores are the benchmark ✓ The Chain of Infection has six links — SPD breaks the chain ✓ Spaulding Classification determines required processing level ✓ Multiple regulatory bodies govern SPD — IFUs are mandatory ✓ Sterile By Design™ = Standardization + Verification + Documentation + Communication

## What's Next

In Module 2, you will learn the practical skills of decontamination — the critical first step in processing reusable medical devices. You will learn:

- Point-of-use treatment and transport

- Manual cleaning techniques

- Ultrasonic cleaning

- Automated washer operation

- Quality verification of the cleaning process

════════════════════════════════════════

Remember: Every instrument you process is going to be used on a human being.

Process every instrument as if it will be used on your own family.

════════════════════════════════════════