SPD TECHNICIAN

TRAINING GUIDE

────────────────────────────────────────

MODULE 4

ASSEMBLY &

TRAY BUILDING

Building Sets with Precision

IN THIS MODULE • The Assembly Workstation • Count Sheets & Pick Lists • Tray Organization Principles • Instrument Protection • Building Common Tray Types • Container Systems • Weight Limits & Sterilization Considerations • Quality Verification Checkpoints

SIPS HEALTHCARE SOLUTIONS

Sterile By Design™ Curriculum

📚 LEARNING OBJECTIVES

After completing this module, you will be able to: 1. Set up and maintain an organized assembly workstation 2. Use count sheets and pick lists to build accurate instrument sets 3. Apply tray organization principles for optimal sterilization and usability 4. Protect instruments from damage during assembly 5. Assemble common tray types following established procedures 6. Prepare rigid container systems correctly 7. Verify tray weight limits and sterilization compatibility 8. Perform quality verification at each assembly checkpoint

❤️ THE PATIENT CONNECTION When a surgeon opens a tray, they need exactly what they expect — every instrument, in the right place, ready to use. A missing instrument can delay surgery. A damaged instrument can fail at a critical moment. Your precision in assembly directly impacts patient care in the operating room.

# 4.1 The Assembly Workstation

Assembly (also called "prep and pack") takes place in the clean area of SPD. This area has specific requirements to maintain the cleanliness of instruments before sterilization.

## Clean Area Requirements

### Environmental Controls

- Temperature: Controlled, typically 68-73°F (20-23°C)

- Humidity: Controlled, typically 30-60% relative humidity

- Air flow: Positive pressure relative to decontam; filtered air

- Traffic: Restricted access; minimize movement through area

### Attire Requirements

- Clean scrubs: Facility-provided; changed if visibly soiled

- Hair cover: Bouffant cap covering all hair

- Clean gloves: Lint-free; changed frequently

- No jewelry: Rings, bracelets, watches removed (small earrings may be permitted)

- Short, clean nails: No artificial nails

⚠️ The Barrier There is a physical and procedural barrier between decontam (dirty) and assembly (clean). You cannot move directly from decontam to assembly without: • Removing decontam PPE • Performing hand hygiene • Changing to clean attire This barrier protects the cleanliness of processed instruments.

## Workstation Setup

An organized workstation promotes efficiency and accuracy.

### Essential Workstation Items

| ITEM | PURPOSE |
| --- | --- |
| Clean work surface | Lint-free, cleaned between uses |
| Lighted magnifier | Final inspection of instruments during assembly |
| Count sheets/pick lists | Reference for tray contents |
| Tip protectors | Protect delicate instrument tips |
| Instrument stringers/racks | Organize and separate instruments in trays |
| Chemical indicators (internal) | Placed inside packs to verify sterilization exposure |
| Labels/markers | Identify trays, document assembly |
| Towels/absorbent liners | Line tray bottoms; absorb moisture |

### Workstation Hygiene

- Clean work surface between each tray or set

- Change gloves frequently and whenever contamination suspected

- Don't eat, drink, or store personal items in the work area

- Report any spills or contamination immediately

# 4.2 Count Sheets & Pick Lists

Count sheets (also called pick lists, tray lists, or tray contents lists) are the foundation of accurate assembly. They specify exactly which instruments belong in each tray.

## Anatomy of a Count Sheet

A typical count sheet includes:

| ELEMENT | DESCRIPTION | EXAMPLE |
| --- | --- | --- |
| Tray name | Official name of the instrument set | Major Basic Tray, Lap Chole Set |
| Tray ID/number | Unique identifier for tracking | MB-001, LC-042 |
| Instrument list | Each instrument with name and quantity | Mayo Scissors, Straight, 6.75" — Qty: 2 |
| Image (if available) | Photo or illustration of each instrument | [Image] |
| Catalog/reference number | Manufacturer's part number for identification | V. Mueller SU1234 |
| Total count | Total number of instruments in the complete tray | Total: 47 instruments |
| Special instructions | Notes on assembly, sterilization, or handling | Include biological indicator; wrap double-layer |
| Revision date | When the count sheet was last updated | Revised: 10/15/2024 |

## Using Count Sheets

### The SIPS Count Sheet Process

- Obtain current count sheet: Always use the most recent version; check revision date

- Review before starting: Familiarize yourself with contents and any special instructions

- Gather instruments: Collect all instruments needed before beginning assembly

- Verify each instrument: Match each instrument to the count sheet entry

- Check off as you go: Mark each item as it is placed in the tray

- Final count: Count total instruments and compare to expected total

- Document: Sign and date the count sheet per facility protocol

⚠️ Never Substitute Without Authorization If an instrument is missing, damaged, or unavailable: • Do NOT substitute a "similar" instrument without approval • Notify your supervisor • Follow facility protocol for substitutions • Document any approved substitutions Unauthorized substitutions can cause problems in surgery.

## Common Count Sheet Errors

| ERROR | CONSEQUENCE | PREVENTION |
| --- | --- | --- |
| Using outdated count sheet | Wrong instruments in tray; count discrepancy in OR | Always verify revision date; use approved source |
| Miscounting instruments | OR count issues; potential retained instrument | Count twice; use check-off method; verify total |
| Wrong size/type | Instrument unavailable when needed | Read descriptions carefully; verify markings |
| Skipping items | Missing instruments in tray | Work systematically; check off each item |
| Unauthorized substitution | Wrong instrument for procedure; surgeon frustration | Always get approval; document substitutions |

# 4.3 Tray Organization Principles

How you arrange instruments in a tray affects sterilization effectiveness, instrument protection, and usability in the OR.

## Goals of Tray Organization

- Sterilization effectiveness: Steam/sterilant must reach all surfaces

- Instrument protection: Prevent damage during sterilization and handling

- Ease of use: OR staff can quickly find and identify instruments

- Accurate counting: Organized layout facilitates counting before and after surgery

- Drying: Allow for drainage and drying

## General Organization Rules

### Instrument Position

- Hinged instruments open: All box locks must be unlocked to allow steam penetration

- Single layer when possible: Avoid stacking instruments directly on top of each other

- Heaviest on bottom: Place heavier instruments at the bottom of the tray

- Delicate items protected: Use tip protectors; position away from heavy items

- Similar items together: Group like instruments for easy identification

### Orientation

- Concave surfaces down: Allows water to drain (bowls, basins, kidney dishes)

- Working ends accessible: Jaws, tips visible for inspection

- Handles toward edge: Easy to grasp; consistent positioning

- Sharp ends protected: Points directed away from wrap; use protectors

## Using Stringers and Racks

Instrument stringers (also called organizing racks, pins, or holders) keep instruments organized and separated within trays.

### Benefits of Stringers

- Keep instruments in designated positions

- Prevent instruments from shifting during handling

- Hold box locks open during sterilization

- Facilitate counting (each position = one instrument)

- Protect instruments from contact damage

### Types of Organizing Systems

| TYPE | DESCRIPTION | BEST FOR |
| --- | --- | --- |
| Pin-type stringers | Metal pins on a base; instruments hang from pins | Hinged instruments (hemostats, clamps) |
| Rail/rack systems | Instruments slide into slots on a rail | Scissors, forceps, varied instrument sizes |
| Silicone mats/holders | Flexible mats with molded positions | Delicate instruments; custom configurations |
| Finger mats | Raised ridges that separate instruments | General instrument separation |
| 🎯 Box Locks Must Be Open This is critical and worth repeating: ALL hinged instruments must have box locks in the OPEN (unlocked) position. Closed box locks prevent steam from reaching the crevices where blood and tissue can hide. If an instrument won't stay open on a stringer, use a rack or position that keeps it open. |

# 4.4 Instrument Protection

Instruments are precision tools — and they're expensive. Proper handling and protection during assembly prevents damage that can make instruments unusable.

## Types of Damage to Prevent

| DAMAGE TYPE | CAUSE | PREVENTION |
| --- | --- | --- |
| Tip damage | Tips striking tray or other instruments | Use tip protectors; handle carefully; proper positioning |
| Edge damage (cutting instruments) | Edges contacting metal surfaces | Blade protectors; separate from other instruments |
| Bending | Weight of other instruments; improper handling | Proper positioning; don't overload trays |
| Surface scratching | Metal-to-metal contact during handling | Use stringers/racks; handle one at a time |
| Insulation damage | Pressure, sharp edges, improper handling | Separate insulated instruments; use protectors |

## Protective Devices

### Tip Protectors

Small caps or covers that fit over delicate instrument tips to prevent damage.

- Use on all delicate, pointed, or sharp-tipped instruments

- Must be appropriate size for the instrument

- Must be sterilization-compatible (vented or perforated)

- Must not prevent steam penetration

### Blade Guards

Protective covers for cutting edges.

- Use on osteotomes, chisels, and other edged instruments

- Must be secure but easy to remove

- Must allow sterilization

### Foam or Silicone Padding

Placed between instruments or lining tray bottoms.

- Reduces metal-to-metal contact

- Must be lint-free and sterilization-compatible

- Check manufacturer guidelines for use

⚠️ Protective Devices Must Not Impede Sterilization Any protective device must allow steam or sterilant to reach instrument surfaces. Solid caps or covers can prevent sterilization and must not be used. Only use protectors approved for sterilization.

## Handling Principles

- One at a time: Handle instruments individually, not in handfuls

- By the body: Grasp instruments by the shaft or body, not delicate tips

- Gently: Place instruments in trays — don't drop or toss them

- Inspect as you go: Final inspection opportunity as you handle each instrument

# 4.5 Building Common Tray Types

Different tray types have specific organization patterns. You'll learn your facility's specific trays over time, but these general principles apply across most tray types.

## General Surgical Tray (Major Basic)

The workhorse tray containing fundamental surgical instruments used in most procedures.

### Typical Organization

- Bottom layer: Absorbent towel or liner for drainage

- Retractors: Larger, heavier items positioned first

- Clamps/hemostats: On stringers, organized by size

- Scissors: Grouped together; tips protected

- Needle holders: Positioned for easy identification

- Forceps: Tissue forceps grouped together

- Specialty items: Positioned per count sheet

- Internal indicator: Placed in center or per facility protocol

## Specialty Trays

### Orthopedic Trays

Contain heavy instruments; often require specific racks or holders.

- Heavy instruments (retractors, bone holders) on bottom

- Power equipment accessories in designated positions

- Implant trays have specific organization requirements

- Weight limits especially important

### Laparoscopic Trays

Contain delicate, precision instruments with lumens.

- Instruments often have long, narrow shafts — protect from bending

- Ensure lumens are open for sterilization

- Delicate optics require careful handling

- Insulated instruments need insulation protection

### Microsurgical/Ophthalmic Trays

Contain extremely delicate instruments.

- Usually organized in specialized holders

- Every instrument needs tip protection

- Handle with extreme care

- Often assembled in specific sequence

💡 Follow Your Facility's Count Sheets Every facility organizes trays slightly differently. The count sheet for your facility is your authoritative guide. When in doubt, ask your preceptor or reference the specific count sheet.

# 4.6 Container Systems

Rigid sterilization containers (also called instrument containers or container systems) are reusable metal or plastic containers used instead of or in addition to wrapping.

## Container Components

| COMPONENT | FUNCTION | INSPECTION POINTS |
| --- | --- | --- |
| Container body (basket) | Holds instruments; may be perforated | Clean; no dents, cracks, or warping |
| Lid | Covers container; contains filter or valve | Latches function; seals properly; no damage |
| Filter (disposable or reusable) | Allows steam in/out while maintaining sterility | Correct filter type; properly seated; not expired |
| Filter retention plate | Holds filter in place | Clean; properly positioned; secure |
| Gasket/seal | Creates seal between lid and container | Intact; pliable; properly seated |
| Latches/locks | Secure lid to container | Function properly; lock securely; release easily |
| Tamper-evident device | Indicates if container was opened | Correct type; applied per manufacturer instructions |

## Container Assembly Process

- Inspect container: Check all components for damage, cleanliness, function

- Check gaskets: Verify seals are intact and properly seated

- Install filter: Place correct filter type; secure with retention plate

- Load instruments: Place instruments per count sheet; use organizing racks

- Place internal indicator: Position chemical indicator inside container

- Close and latch: Secure lid; verify all latches engage

- Apply external indicator: Attach tamper-evident seal or lock

- Label: Apply labels per facility protocol

⚠️ Filter Is Critical The filter is what allows steam to enter and exit while keeping bacteria out. • Wrong filter type = sterilization failure or sterility breach • Damaged filter = compromised sterility • Missing filter = container cannot maintain sterility Always verify the correct filter is installed before closing the container.

## Container Advantages and Considerations

### Advantages

- Better instrument protection than wraps

- Environmental benefits (reusable)

- Consistent presentation

- Long shelf life when stored properly

### Considerations

- Must follow manufacturer IFU for assembly and sterilization

- Components must be inspected and maintained

- Specific loading guidelines must be followed

- Weight limits apply

# 4.7 Weight Limits & Sterilization Considerations

Tray weight affects sterilization effectiveness, safe handling, and instrument drying. Following weight limits is not optional.

## Weight Limits

### AAMI Recommendations

| CONFIGURATION | MAXIMUM WEIGHT |
| --- | --- |
| Wrapped tray (set) | 25 lbs (11.3 kg) maximum |
| Rigid container (per manufacturer IFU) | Typically 25 lbs (11.3 kg); check specific container IFU |
| 📌 Why Weight Matters Overweight trays cause multiple problems: • Steam may not penetrate adequately • Drying is inadequate (wet packs) • Increased risk of handling injuries • Instrument damage from compression • Container or wrap failure |

## Sterilization Compatibility

Not everything can go in the same load. Consider compatibility when assembling trays.

### Sterilization Method Compatibility

| ITEM TYPE | STEAM (AUTOCLAVE) | LOW-TEMP (H2O2/EO) |
| --- | --- | --- |
| Stainless steel instruments | ✓ Yes | ✓ Yes (usually not needed) |
| Heat-sensitive plastics | ✗ No | ✓ Yes (check IFU) |
| Power equipment | Usually no (check IFU) | ✓ Yes (check IFU) |
| Flexible scopes | ✗ No | ✓ Yes (check IFU) |
| Batteries | ✗ No — remove before sterilization | ✗ No — remove before sterilization |

Always check the manufacturer IFU for each device to confirm sterilization compatibility.

## Tray Size and Configuration

### Tray Selection

- Use smallest tray that accommodates instruments properly

- Instruments should fit without crowding or forcing

- Leave space for steam circulation

- Consider sterilizer shelf capacity

### Multiple Levels

Some trays have multiple tiers or levels:

- Heaviest instruments on bottom level

- Lighter/delicate items on upper levels

- Ensure steam can reach all levels

- Follow count sheet for level assignments

# 4.8 Quality Verification Checkpoints

Quality verification during assembly catches errors before they reach the OR. Use these checkpoints for every tray you build.

## SIPS Assembly Verification Checkpoints

| CHECKPOINT | VERIFICATION |
| --- | --- |
| ☐ 1. Count Sheet | Using current version; all items present |
| ☐ 2. Instrument Cleanliness | Each instrument visually clean; no soil, staining, debris |
| ☐ 3. Instrument Function | All instruments function properly; no damage |
| ☐ 4. Box Locks Open | All hinged instruments in unlocked position |
| ☐ 5. Organization | Instruments positioned per count sheet; properly protected |
| ☐ 6. Weight | Within weight limit (≤25 lbs) |
| ☐ 7. Internal Indicator | Chemical indicator placed inside pack/container |
| ☐ 8. Final Count | Total count matches count sheet |
| ☐ 9. Container/Wrap Ready | Container assembled correctly OR ready for wrapping |
| ☐ 10. Documentation | Count sheet signed/initialed; tracking scan complete |

## Common Assembly Errors

| ERROR | CONSEQUENCE | PREVENTION |
| --- | --- | --- |
| Wrong count | OR count discrepancy; potential retained instrument | Count twice; use checkmarks; verify total |
| Box locks closed | Sterilization failure; potential infection | Final check of every hinged instrument |
| Missing indicator | Cannot verify sterilization; tray unusable | Include indicator as standard step |
| Overweight | Wet pack; inadequate sterilization | Weigh trays; know limits |
| Wrong instruments | Surgeon doesn't have what's needed | Verify each instrument against count sheet |
| Damaged instruments included | Instrument fails during surgery | Inspect every instrument during assembly |

🎯 Your Signature Means Something When you sign or initial a count sheet, you are attesting that: • All instruments are present and accounted for • All instruments are clean and functional • The tray is correctly assembled • You would trust this tray for your own surgery Never sign off on a tray you haven't personally verified.

# Knowledge Check: Module 4

Answer these questions to verify your understanding of this module. Discuss any questions with your preceptor.

### Section 4.1 — Assembly Workstation

- What attire is required in the assembly (clean) area?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 4.2 — Count Sheets

- What should you do if an instrument listed on the count sheet is missing or damaged?

_____________________________________________________________________________

_____________________________________________________________________________

- Why is it important to use the most current version of a count sheet?

_____________________________________________________________________________

### Section 4.3 — Tray Organization

- Why must all hinged instruments have box locks in the open position?

_____________________________________________________________________________

_____________________________________________________________________________

- What is the purpose of using instrument stringers/racks?

_____________________________________________________________________________

### Section 4.4 — Instrument Protection

- List three types of damage that can occur during assembly and how to prevent each:

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

### Section 4.6 — Container Systems

- What is the function of the filter in a rigid container system?

_____________________________________________________________________________

_____________________________________________________________________________

- What components should you inspect before using a container?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 4.7 — Weight Limits

- What is the maximum recommended weight for a wrapped instrument tray?

_____________________________________________________________________________

- What problems can occur if a tray exceeds the weight limit?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 4.8 — Quality Verification

- List five quality verification checkpoints you should complete for every tray:

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

4. ___________________________________________________________________________

5. ___________________________________________________________________________

# Skills Validation: Module 4

Your preceptor will observe you performing the following skills. Each skill must be demonstrated correctly.

| SKILL | MET | NOT MET | DATE | INITIALS |
| --- | --- | --- | --- | --- |
| Wears proper attire for assembly area | ☐ | ☐ |  |  |
| Sets up and maintains organized workstation | ☐ | ☐ |  |  |
| Uses count sheet correctly; verifies current version | ☐ | ☐ |  |  |
| Assembles tray per count sheet with correct count | ☐ | ☐ |  |  |
| Organizes instruments properly (position, orientation) | ☐ | ☐ |  |  |
| Ensures all box locks are open | ☐ | ☐ |  |  |
| Uses instrument protectors appropriately | ☐ | ☐ |  |  |
| Assembles rigid container correctly | ☐ | ☐ |  |  |
| Verifies weight is within limits | ☐ | ☐ |  |  |
| Places internal chemical indicator | ☐ | ☐ |  |  |
| Completes all verification checkpoints | ☐ | ☐ |  |  |
| Documents assembly correctly | ☐ | ☐ |  |  |

### Module Completion

| Learner Name: | ________________________________________ |
| --- | --- |
| Learner Signature: | ________________________________________ |
| Date: | ________________________________________ |
| Preceptor Name: | ________________________________________ |
| Preceptor Signature: | ________________________________________ |
| Date: | ________________________________________ |

# Module 4 Summary

✓ Key Takeaways ✓ Assembly area requires clean attire and proper hand hygiene ✓ Always use the current version of count sheets ✓ Never substitute instruments without authorization ✓ All hinged instruments must have box locks OPEN ✓ Use stringers and protectors to organize and protect instruments ✓ Container filters are critical — verify correct type and installation ✓ Maximum wrapped tray weight: 25 lbs ✓ Complete all verification checkpoints for every tray ✓ Your signature = your attestation that the tray is correct

## What's Next

In Module 5, you will learn packaging and wrapping — how to prepare instrument sets for sterilization. You will learn:

- Types of sterilization wrap

- Wrapping techniques (envelope, square fold)

- Peel pack preparation

- Labeling requirements

- Package integrity

════════════════════════════════════════

Precision in assembly means confidence in the OR.

Every count. Every instrument. Every time.

════════════════════════════════════════