SPD TECHNICIAN

TRAINING GUIDE

────────────────────────────────────────

MODULE 6

STERILIZATION

The Science of Eliminating Microbial Life

IN THIS MODULE • Principles of Sterilization • Steam Sterilization (Autoclaving) • Low-Temperature Sterilization Methods • Loading Sterilizers • Cycle Parameters & Monitoring • Biological Indicators • Documentation & Record Keeping • Sterilization Failures & Troubleshooting

SIPS HEALTHCARE SOLUTIONS

Sterile By Design™ Curriculum

📚 LEARNING OBJECTIVES

After completing this module, you will be able to: 1. Explain the principles of sterilization and how sterilization differs from disinfection 2. Describe how steam sterilization works and identify critical cycle parameters 3. Identify low-temperature sterilization methods and their appropriate applications 4. Load sterilizers correctly for effective sterilization 5. Monitor cycle parameters and recognize abnormal readings 6. Explain the purpose and use of biological indicators 7. Complete required sterilization documentation accurately 8. Recognize sterilization failures and respond appropriately

❤️ THE PATIENT CONNECTION Sterilization is what makes surgical instruments safe for use inside the human body. A single surviving microorganism on a surgical instrument can cause a life-threatening infection. Your understanding of sterilization principles and your attention to process parameters protects every patient.

# 6.1 Principles of Sterilization

Sterilization is the complete elimination or destruction of all forms of microbial life, including highly resistant bacterial spores.

## Sterilization vs. Disinfection

|  | STERILIZATION | HIGH-LEVEL DISINFECTION |
| --- | --- | --- |
| Definition | Destruction of ALL microbial life including spores | Destruction of all microorganisms EXCEPT high numbers of bacterial spores |
| Use for | Critical items (enter sterile tissue or vascular system) | Semi-critical items (contact mucous membranes) |
| Examples | Surgical instruments, implants | Endoscopes, respiratory equipment |
| Verification | Biological indicators (spore tests) | Chemical indicators, testing protocols |

## How Sterilization Works

All sterilization methods work by disrupting essential cellular functions of microorganisms:

- Heat (steam): Denatures proteins and disrupts cell membranes

- Hydrogen peroxide plasma: Oxidizes cellular components through free radicals

- Ethylene oxide (EO): Alkylates DNA and proteins, preventing reproduction

- Ozone: Oxidizes cellular components

## Factors Affecting Sterilization

| FACTOR | IMPACT |
| --- | --- |
| Bioburden | More microorganisms = longer exposure needed; cleaning reduces bioburden |
| Organic matter | Blood, tissue, and debris protect microorganisms from sterilant |
| Device design | Lumens, crevices, and complex designs require longer exposure |
| Packaging | Must allow sterilant penetration; improper packaging prevents sterilization |
| Loading | Overloading prevents sterilant from reaching all surfaces |
| Cycle parameters | Temperature, time, pressure, and concentration must be correct |

⚠️ Sterilization Cannot Compensate for Poor Cleaning Sterilization works on clean instruments. Organic matter (blood, tissue) shields microorganisms from the sterilant. Biofilm in lumens can survive sterilization. This is why decontamination (Module 2) is so critical.

# 6.2 Steam Sterilization

Steam sterilization (autoclaving) is the most common and preferred method for sterilizing heat- and moisture-stable medical devices. It uses saturated steam under pressure.

## How Steam Sterilization Works

Steam sterilizes through a combination of:

- Heat: High temperature denatures microbial proteins

- Moisture: Steam carries heat more efficiently than dry air; moisture aids protein denaturation

- Pressure: Allows steam to reach temperatures above 212°F (100°C)

- Time: Sufficient exposure ensures all microorganisms are killed

💡 Why Pressure? At normal atmospheric pressure, water boils at 212°F (100°C). Higher pressure raises the boiling point, creating hotter steam: • 15 psi above atmospheric → 250°F (121°C) • 27 psi above atmospheric → 270°F (132°C) Hotter steam = faster sterilization.

## Types of Steam Sterilizers

### Gravity Displacement

Air is displaced by gravity as steam enters from the top. Air (heavier than steam) exits through a drain at the bottom.

- Advantages: Simple design; reliable; good for non-porous items

- Limitations: Slower air removal; not ideal for porous loads or wrapped items

- Typical use: Unwrapped instruments, glassware, liquids, waste

### Prevacuum (Dynamic Air Removal)

Vacuum pumps actively remove air before steam is introduced. Multiple vacuum and steam pulses ensure complete air removal.

- Advantages: Faster cycles; better steam penetration; ideal for wrapped items and porous loads

- Limitations: More complex; requires daily air removal testing (Bowie-Dick)

- Typical use: Wrapped instrument trays, packs, containers, textiles

### Steam Flush-Pressure Pulse (SFPP)

Repeated sequences of steam flush and pressure pulses remove air without deep vacuum.

- Advantages: Good air removal; may be gentler on some items

- Typical use: Similar to prevacuum; follow manufacturer guidance

## Steam Sterilization Parameters

| CYCLE TYPE | TEMPERATURE | EXPOSURE TIME | TYPICAL USE |
| --- | --- | --- | --- |
| Gravity — 250°F | 250°F (121°C) | 30 minutes | Unwrapped instruments, glassware |
| Gravity — 270°F | 270°F (132°C) | 15 minutes | Unwrapped instruments |
| Prevacuum — 270°F | 270°F (132°C) | 4 minutes | Wrapped trays, porous loads |
| Prevacuum — 275°F | 275°F (135°C) | 3 minutes | Wrapped trays, porous loads |
| Extended cycle | 270°F (132°C) | 10+ minutes | Complex devices, per IFU |

Note: Always follow device manufacturer IFU for specific cycle requirements. Some devices require extended exposure times.

# 6.3 Low-Temperature Sterilization

Some medical devices cannot tolerate the heat and moisture of steam sterilization. Low-temperature methods are used for these heat-sensitive items.

## Hydrogen Peroxide Gas Plasma

Uses hydrogen peroxide vapor converted to plasma (ionized gas) by radiofrequency energy.

| CHARACTERISTIC | DETAIL |
| --- | --- |
| Temperature | Low (104-131°F / 40-55°C) |
| Cycle time | 28-75 minutes depending on cycle and system |
| Advantages | No toxic residues; short cycle times; no aeration needed |
| Limitations | Cannot use cellulose (paper/cotton); limited lumen length/diameter; absorptive materials |
| Common use | Heat-sensitive devices, electronics, some scopes, power equipment |
| Brand examples | STERRAD® (H₂O₂ plasma) |

⚠️ No Cellulose in H₂O₂ Plasma Paper, cotton, muslin, and cellulose-based materials absorb hydrogen peroxide and will cause cycle failure. Use only compatible packaging (Tyvek pouches, polypropylene wraps, dedicated containers).

## Vaporized Hydrogen Peroxide (VHP)

Uses hydrogen peroxide vapor without plasma phase. Different from H₂O₂ plasma.

- Temperature: Low (77-122°F / 25-50°C)

- Cycle time: Varies by system and load

- Advantages: Can process longer/narrower lumens than plasma systems

- Brand examples: V-PRO®, STERIS systems

## Ethylene Oxide (EO)

Gas sterilization using ethylene oxide, a chemical that penetrates packaging and device lumens.

| CHARACTERISTIC | DETAIL |
| --- | --- |
| Temperature | Low (99-140°F / 37-60°C) |
| Cycle time | 1-6 hours exposure + 8-12 hours aeration |
| Advantages | Excellent penetration; compatible with many materials; can sterilize long lumens |
| Limitations | Long cycle time; toxic — requires aeration; regulated emissions; flammable |
| Common use | Heat-sensitive devices not compatible with other methods; complex devices |

🛡️ EO Safety Ethylene oxide is toxic, carcinogenic, and flammable. Strict safety protocols apply: • Dedicated ventilated area • Personal protective equipment • Environmental monitoring • Proper aeration before use Never release items before aeration is complete.

## Ozone Sterilization

Uses ozone (O₃) generated from oxygen to sterilize devices.

- Temperature: Low (85-95°F / 30-35°C)

- Cycle time: Approximately 4 hours

- Advantages: Converts back to oxygen; environmentally friendly; no toxic residue

- Limitations: Limited material compatibility; not widely available

# 6.4 Loading Sterilizers

Proper loading is essential for effective sterilization. Incorrect loading is a leading cause of sterilization failures.

## General Loading Principles

- Don't overload: Sterilant must circulate freely around all packages

- Allow for expansion: Packs expand during sterilization; leave room

- Position for drainage: Items should allow condensate to drain

- Separate packages: Packages should not touch each other or chamber walls

- Use appropriate cycle: Match cycle to load contents

## Steam Sterilizer Loading

### Wrapped Packages and Trays

- Place packages on edge (vertically) when possible to allow steam penetration

- Do not stack packages directly on top of each other

- Leave space between packages (at least 1 inch)

- Place heavier packs on bottom shelves

- Paper/plastic pouches positioned paper-to-paper or plastic-to-plastic

### Containers

- Ensure filters/valves are in place and functioning

- Position containers with filter/valve facing up (if specified by manufacturer)

- Allow space around containers

### Mixed Loads

- If mixing porous and non-porous items, use prevacuum cycle

- Non-porous items on bottom, porous items on top

- Never mix incompatible items (e.g., steam-compatible with heat-sensitive)

🎯 Loading Determines Success The best sterilizer running the correct cycle cannot compensate for improper loading. • Overcrowded loads = poor steam penetration = sterilization failure • Touching packages = wet packs • Improper orientation = poor drying Take the time to load correctly.

## Low-Temperature Sterilizer Loading

### Hydrogen Peroxide Systems

- Use ONLY compatible packaging (no cellulose/paper)

- Do not overload — hydrogen peroxide needs to contact all surfaces

- Ensure lumened devices meet length/diameter specifications

- Follow manufacturer loading configurations exactly

### EO Sterilizers

- Allow adequate space for gas circulation

- Do not overload

- Position packages for gas penetration

- Follow facility and manufacturer protocols

# 6.5 Cycle Parameters & Monitoring

Every sterilization cycle must be monitored to verify that correct parameters were achieved.

## Critical Cycle Parameters

### Steam Sterilization

| PARAMETER | WHY IT MATTERS | MONITORING |
| --- | --- | --- |
| Temperature | Must reach target temp for sterilization to occur | Printout/digital display shows temperature throughout cycle |
| Time | Exposure time must be sufficient at target temperature | Printout shows exposure time |
| Pressure | Indicates proper steam generation and air removal | Printout/gauge shows pressure |
| Air removal (prevac) | Air prevents steam contact; must be removed | Bowie-Dick test daily; vacuum readings |
| Drying | Wet packs are not sterile | Visual inspection; dry time on printout |

### Low-Temperature Sterilization

- H₂O₂ systems: Concentration, pressure, plasma generation, cycle completion

- EO: Gas concentration, humidity, temperature, exposure time, aeration

## Monitoring Methods

### Physical Monitors

Built-in gauges, digital displays, and printouts that record cycle parameters.

- Review printout after every cycle

- Verify temperature reached target

- Verify exposure time was complete

- Check for any alarms or error messages

### Chemical Indicators

Provide visual confirmation of exposure to sterilization conditions.

- External (Class 1): Distinguishes processed from unprocessed items

- Internal (Class 4, 5, 6): Verifies sterilant reached inside the package

### Biological Indicators (BIs)

The gold standard for verifying sterilization — contains live spores that should be killed by proper sterilization.

BIs are covered in detail in Section 6.6.

## Bowie-Dick Test

The Bowie-Dick test is performed daily in prevacuum steam sterilizers to verify adequate air removal.

### When to Perform

- Daily, before the first processed load

- After sterilizer repairs or maintenance

- After sterilizer malfunction

### How It Works

- Commercial test pack placed in empty sterilizer

- Run specific Bowie-Dick cycle

- Examine test sheet for uniform color change

- Pass: Uniform color change across entire sheet

- Fail: Uneven color, lighter center, or incomplete change indicates air leak or inadequate air removal

⚠️ Failed Bowie-Dick = No Processing If the Bowie-Dick test fails: • Do NOT use the sterilizer for patient care items • Report immediately to supervisor/biomedical • Investigate and repeat test after issue resolved

# 6.6 Biological Indicators

Biological indicators (BIs) contain live bacterial spores — the most resistant form of microbial life. If the BI spores are killed, you can be confident the sterilization process is effective.

## How BIs Work

- BI contains a known population of highly resistant bacterial spores

- BI is placed in the sterilizer with the load (in a challenge location)

- After sterilization, BI is incubated to see if any spores survived

- Negative (no growth) = PASS: Spores were killed; sterilization effective

- Positive (growth) = FAIL: Spores survived; sterilization failure

## Types of BIs

| STERILIZATION METHOD | SPORE TYPE | INCUBATION |
| --- | --- | --- |
| Steam | Geobacillus stearothermophilus | 55-60°C (131-140°F); 24-48 hours (rapid: 1-3 hours) |
| H₂O₂ Plasma/VHP | Geobacillus stearothermophilus | 55-60°C; 24 hours (rapid: 30 min - 1 hour) |
| Ethylene Oxide | Bacillus atrophaeus | 35-37°C (95-99°F); 48 hours (rapid: 4 hours) |

## When to Use BIs

| SITUATION | BI REQUIREMENT |
| --- | --- |
| Routine monitoring | At least daily, preferably every load |
| Every load containing implants | REQUIRED — do not release implants until BI results negative |
| After sterilizer installation | Multiple test cycles before routine use |
| After sterilizer repair | Three consecutive negative BIs before returning to service |
| After sterilization failure | Three consecutive negative BIs after issue resolved |

## BI Placement

Place BI in the location most challenging for sterilant to reach:

- Steam (gravity): Center of the load, lower shelf

- Steam (prevac): Over the drain, front/bottom of the load

- H₂O₂/EO: Per manufacturer instructions; often center of load

- In a pack: Place inside a pack or container for most accurate challenge

## Responding to Positive BI

⚠️ Positive BI = Sterilization Failure If a BI comes back positive (growth detected): 1. IMMEDIATELY notify supervisor 2. Quarantine ALL items from that load (if not already used) 3. Recall items if possible and clinically appropriate 4. Investigate cause (review parameters, loading, equipment) 5. Do not use sterilizer until issue resolved 6. Run three consecutive negative BIs before returning to service 7. Document thoroughly

# 6.7 Documentation & Record Keeping

Complete documentation provides evidence that sterilization was performed correctly and enables traceability if problems occur.

## What Must Be Documented

| DOCUMENT | CONTENT | RETENTION |
| --- | --- | --- |
| Cycle printout/record | Date, time, sterilizer ID, cycle number, parameters achieved, operator | Per facility policy (often 3-10 years) |
| Load contents | What items were in each load | Per facility policy |
| BI results | Date, sterilizer, load, BI lot number, result, technician | Per facility policy |
| Bowie-Dick results | Date, sterilizer, result, technician | Per facility policy |
| Maintenance records | PM schedule, repairs, calibration | Per facility policy and regulations |
| Failures/incidents | Any cycle failures, positive BIs, equipment malfunctions | Per facility policy |

## Reviewing Cycle Printouts

After every cycle, review the printout to verify:

- Cycle completed: No aborts or errors

- Temperature reached target: Check exposure phase temperature

- Exposure time complete: Full time at target temperature

- Pressure appropriate: Consistent with normal readings

- No alarms: No error codes or warnings

- Drying completed: For steam sterilizers

If ANY parameter is outside acceptable range — do NOT release the load. Investigate and reprocess if necessary.

## Traceability

Traceability allows you to track exactly which sterilization cycle processed a specific item. This is essential for recalls and investigations.

- Each package labeled: Sterilizer ID, load/cycle number, date

- Implants: Lot numbers tracked to specific patients

- Tracking systems: Many facilities use barcode or RFID tracking

🎯 If It Isn't Documented, It Didn't Happen In healthcare, documentation is essential for quality and safety. • Complete records protect patients (enable recalls) • Complete records protect you (demonstrate proper practice) • Complete records protect the facility (demonstrate compliance)

# 6.8 Sterilization Failures & Troubleshooting

Understanding what causes sterilization failures helps you prevent them and respond appropriately when they occur.

## Types of Failures

| FAILURE TYPE | DESCRIPTION | INDICATORS |
| --- | --- | --- |
| Cycle abort | Sterilizer stops before cycle completes | Error code; incomplete printout; alarm |
| Parameter failure | Target parameters not achieved | Printout shows incorrect temp, time, or pressure |
| Wet pack | Moisture on or in packages after sterilization | Visible moisture; damp packaging |
| Chemical indicator failure | CI does not show complete color change | Visual inspection of indicator |
| Biological indicator positive | Spores survived sterilization | BI shows growth after incubation |
| Bowie-Dick failure | Inadequate air removal (prevac) | Uneven color change on test sheet |

## Common Causes of Failures

| CAUSE | RESULT | PREVENTION |
| --- | --- | --- |
| Overloading | Poor steam penetration; wet packs | Follow loading guidelines; allow space |
| Improper loading | Steam cannot reach all surfaces | Position packages correctly; paper side to paper side |
| Incorrect cycle selection | Parameters don't match load requirements | Verify cycle matches contents |
| Equipment malfunction | Parameters not achieved; cycle aborts | Regular maintenance; daily testing |
| Packaging error | Sterilant cannot penetrate; compromised barrier | Proper wrapping technique; correct materials |
| Residual air (steam) | Air prevents steam contact | Daily Bowie-Dick test; proper loading |
| Incompatible packaging (H₂O₂) | Cycle cancels or fails | Use only compatible materials; no cellulose |

## Response to Failures

### Immediate Actions

- Do NOT release items: Quarantine the entire load

- Notify supervisor: Report the failure immediately

- Document: Record all details of the failure

- Investigate: Review parameters, loading, equipment

- Reprocess: Repackage and re-sterilize affected items

### Before Returning to Service

- Identify and correct the cause

- For BI positive or major failure: three consecutive negative BIs

- For Bowie-Dick failure: pass Bowie-Dick before processing loads

- Document all corrective actions

❤️ When in Doubt, Don't Release If you are unsure whether a cycle was successful — do not release the load. It is always safer to reprocess than to use potentially non-sterile items. Consult your supervisor if you have any questions about cycle results.

# Knowledge Check: Module 6

Answer these questions to verify your understanding of this module. Discuss any questions with your preceptor.

### Section 6.1 — Principles

- What is the difference between sterilization and high-level disinfection?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 6.2 — Steam Sterilization

- What are the four factors that make steam sterilization effective?

1. ____________________________  2. ____________________________

3. ____________________________  4. ____________________________

- What is the difference between gravity and prevacuum steam sterilizers?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 6.3 — Low-Temperature Sterilization

- Why can't cellulose materials (paper, cotton) be used in H₂O₂ plasma sterilization?

_____________________________________________________________________________

### Section 6.4 — Loading

- List three principles of proper sterilizer loading:

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

### Section 6.5 — Monitoring

- What is the purpose of the Bowie-Dick test?

_____________________________________________________________________________

- What should you check when reviewing a cycle printout?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 6.6 — Biological Indicators

- What does a positive biological indicator result mean, and what should you do?

_____________________________________________________________________________

_____________________________________________________________________________

- When must you wait for BI results before releasing items?

_____________________________________________________________________________

### Section 6.8 — Failures

- List three common causes of sterilization failures:

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

# Skills Validation: Module 6

Your preceptor will observe you performing the following skills. Each skill must be demonstrated correctly.

| SKILL | MET | NOT MET | DATE | INITIALS |
| --- | --- | --- | --- | --- |
| Identifies sterilization methods and appropriate uses | ☐ | ☐ |  |  |
| Selects correct cycle for load contents | ☐ | ☐ |  |  |
| Loads steam sterilizer correctly | ☐ | ☐ |  |  |
| Performs Bowie-Dick test correctly | ☐ | ☐ |  |  |
| Reviews and interprets cycle printout | ☐ | ☐ |  |  |
| Places biological indicator in correct location | ☐ | ☐ |  |  |
| Incubates and reads biological indicator | ☐ | ☐ |  |  |
| Identifies wet pack and responds correctly | ☐ | ☐ |  |  |
| Documents sterilization records accurately | ☐ | ☐ |  |  |
| Responds appropriately to sterilization failure scenario | ☐ | ☐ |  |  |
| Loads low-temperature sterilizer correctly (if applicable) | ☐ | ☐ |  |  |

### Module Completion

| Learner Name: | ________________________________________ |
| --- | --- |
| Learner Signature: | ________________________________________ |
| Date: | ________________________________________ |
| Preceptor Name: | ________________________________________ |
| Preceptor Signature: | ________________________________________ |
| Date: | ________________________________________ |

# Module 6 Summary

✓ Key Takeaways ✓ Sterilization eliminates ALL microbial life including spores ✓ Steam sterilization uses heat, moisture, pressure, and time ✓ Low-temp methods (H₂O₂, EO) are for heat-sensitive items ✓ Proper loading is essential — overloading causes failures ✓ Bowie-Dick test verifies air removal in prevacuum sterilizers ✓ Chemical indicators show exposure; biological indicators prove kill ✓ Implant loads require BI results before release ✓ Positive BI = sterilization failure — quarantine and investigate ✓ Document everything — if it isn't documented, it didn't happen ✓ When in doubt, don't release — reprocess is always safer

## What's Next

In Module 7, you will learn sterile storage and distribution — how to properly store sterile items and get them to the point of use. You will learn:

- Sterile storage area requirements

- Shelf life and event-related sterility

- Stock rotation and inventory management

- Distribution and transport

- Case cart preparation

════════════════════════════════════════

Sterilization is the science. Verification is the proof.

Every cycle. Every load. Every time.

════════════════════════════════════════