SPD TECHNICIAN

TRAINING GUIDE

────────────────────────────────────────

MODULE 7

STERILE STORAGE

& DISTRIBUTION

Protecting Sterility Until Point of Use

IN THIS MODULE • Sterile Storage Area Requirements • Event-Related Sterility • Package Inspection Before Storage • Stock Rotation & Inventory Management • Distribution & Transport • Case Cart Preparation • Handling Recalls & Expired Items

SIPS HEALTHCARE SOLUTIONS

Sterile By Design™ Curriculum

📚 LEARNING OBJECTIVES

After completing this module, you will be able to: 1. Describe the environmental requirements for sterile storage areas 2. Explain event-related sterility and factors that compromise sterility 3. Inspect packages for integrity before placing in storage 4. Apply stock rotation principles (FIFO) to maintain inventory 5. Transport sterile items safely while maintaining sterility 6. Prepare case carts accurately according to pick lists 7. Identify and properly handle recalled or expired items 8. Maintain documentation for storage and distribution activities

❤️ THE PATIENT CONNECTION All the work of cleaning, inspecting, assembling, packaging, and sterilizing can be undone in seconds by improper storage or handling. A compromised package means a compromised instrument — and a risk to the patient. Your care in storage and distribution is the final protection before items reach the OR.

# 7.1 Sterile Storage Area Requirements

Sterile storage areas must be designed and maintained to protect the sterility of processed items until they are needed for patient care.

## Environmental Controls

| PARAMETER | REQUIREMENT | WHY IT MATTERS |
| --- | --- | --- |
| Temperature | ≤75°F (24°C) | High temperatures can degrade packaging materials and compromise seals |
| Humidity | ≤70% relative humidity | High humidity promotes microbial growth and can weaken packaging |
| Air quality | Positive pressure; filtered air; limited traffic | Prevents contamination from less clean areas |
| Lighting | Adequate for inspection; avoid direct sunlight | UV light can degrade some packaging; need visibility for inspection |

## Physical Requirements

### Shelving

- Bottom shelf: At least 8-10 inches off the floor

- Top shelf: At least 18 inches from ceiling/sprinklers

- Distance from walls: At least 2 inches from outside walls

- Material: Smooth, cleanable surfaces; no exposed wood or cardboard

- Design: Open wire shelving preferred for air circulation

### Location and Access

- Separate from decontamination and soiled areas

- Limited traffic — only authorized personnel

- No through-traffic to other areas

- No water pipes or drains overhead (if possible)

- Away from external doors and loading docks

💡 Why the Height Requirements? • Bottom shelf 8-10" off floor: Protects from floor cleaning, flooding, contamination • Top shelf 18" from ceiling: Allows air circulation; prevents contamination from sprinklers; access for cleaning • 2" from walls: Air circulation; prevents moisture accumulation; access for cleaning

## Restricted Items in Sterile Storage

The following should NOT be in sterile storage areas:

- Cardboard boxes (harbor dust, pests, moisture)

- Corrugated shipping containers

- Personal items (food, drinks, bags)

- Soiled or contaminated items

- Trash or recycling

- Equipment not related to sterile storage

# 7.2 Event-Related Sterility

Event-related sterility means that properly packaged and stored items remain sterile until an event occurs that compromises the packaging. Time alone does not make an item non-sterile.

## Events That Compromise Sterility

| EVENT | EXAMPLES | ACTION |
| --- | --- | --- |
| Physical damage | Tears, holes, punctures, crushed packages | Remove from service; reprocess |
| Moisture | Wet packages, water stains, humidity damage | Remove from service; reprocess |
| Seal compromise | Broken seals, open pouches, loose tape | Remove from service; reprocess |
| Contamination | Dropped on floor, contact with contaminated surface | Remove from service; reprocess |
| Improper handling | Excessive handling, compression, stacking damage | Inspect; reprocess if integrity questioned |
| Environmental exposure | Dust accumulation, pest exposure, chemical exposure | Remove from service; reprocess |

## Shelf Life Policies

While event-related sterility is the current standard, facilities may still assign shelf life dates for practical reasons:

- Inventory management: Ensures stock rotation

- Inspection prompts: Items reviewed periodically

- Packaging degradation: Some materials degrade over time

- Manufacturer requirements: Some IFUs specify maximum shelf life

🎯 Time vs. Events A package with an "expiration date" is still sterile if: • The package is intact • No contaminating events have occurred • Storage conditions have been proper However, always follow your facility's policy on shelf life. When dates are assigned, honor them.

## Factors Affecting Shelf Life

| FACTOR | SHORTER SHELF LIFE | LONGER SHELF LIFE |
| --- | --- | --- |
| Packaging type | Paper/plastic pouches; single wrap | Rigid containers; double wrap |
| Storage location | Open shelving; high traffic areas | Closed cabinets; restricted access |
| Environmental conditions | Variable temp/humidity; poor air quality | Controlled environment; filtered air |
| Handling frequency | Frequently moved; high handling | Minimal handling; dedicated location |

# 7.3 Package Inspection Before Storage

Every package must be inspected before it is placed in sterile storage. This is your last chance to catch problems before items are available for patient care.

## Post-Sterilization Inspection Checklist

| INSPECTION POINT | WHAT TO CHECK | IF FAILED |
| --- | --- | --- |
| ☐ External indicator | Color change complete and appropriate | Do not store; investigate; reprocess |
| ☐ Package integrity | No tears, holes, punctures, or damage | Do not store; repackage and reprocess |
| ☐ Seal integrity | All seals intact; tape secure | Do not store; repackage and reprocess |
| ☐ Dryness | Package completely dry inside and out | Do not store; WET PACK — reprocess |
| ☐ Label | Complete, legible, accurate | Do not store without proper identification |
| ☐ Container locks (if applicable) | Latches secure; tamper seals intact | Do not store; inspect contents; reprocess if needed |

## Handling Items That Fail Inspection

- Segregate immediately: Do not place in sterile storage

- Identify the problem: Document what is wrong

- Determine cause: Was it packaging, loading, or equipment issue?

- Take corrective action: Repackage and reprocess; report patterns

- Document: Record the failure and action taken

⚠️ Wet Packs Are Never Acceptable A wet pack is considered contaminated because moisture provides a pathway for microorganisms. If ANY moisture is present: • Do NOT store • Do NOT use • Completely repackage and reprocess • Investigate and document the cause

# 7.4 Stock Rotation & Inventory Management

Proper inventory management ensures sterile items are used efficiently and safely, minimizing waste and ensuring availability.

## FIFO: First In, First Out

FIFO (First In, First Out) is the standard method for stock rotation in sterile storage.

### How FIFO Works

- Oldest items (earliest sterilization date) are used first

- New items are placed behind or below existing stock

- Items are pulled from the front or top

- This ensures stock is used before potential expiration

### Implementing FIFO

- Check dates when putting away new stock

- Move older items forward

- Place newer items in the back

- Never place new stock in front of old stock

💡 FIFO Protects Patients and Budgets FIFO ensures: • Items are used before they expire or are questioned • Reduced waste from expired stock • Consistent inventory turnover • Items don't sit unused for extended periods

## Inventory Organization

### Organization Principles

- Like with like: Group similar items together

- Logical arrangement: Organize by specialty, procedure type, or use frequency

- Consistent locations: Items always in the same place

- Clear labeling: Shelves and bins clearly marked

- Accessibility: Frequently used items easily accessible

### Par Levels

Par levels are minimum and maximum quantities to maintain for each item.

- Minimum (reorder point): When to order more stock

- Maximum: Don't exceed this quantity

- Purpose: Ensures availability without overstocking

## Inventory Counts

| COUNT TYPE | FREQUENCY | PURPOSE |
| --- | --- | --- |
| Perpetual | Ongoing; updated with each transaction | Real-time tracking; used in automated systems |
| Cycle counts | Periodic; rotating sections | Verify accuracy without full inventory shutdown |
| Physical inventory | Annually or as required | Complete verification; reconciliation |

# 7.5 Distribution & Transport

Sterile items must be transported safely to maintain sterility from storage to the point of use.

## Transport Principles

- Protect from damage: Handle carefully; don't drop, crush, or compress

- Protect from moisture: Keep dry; cover during transport

- Protect from contamination: Use covered carts; avoid contaminated areas

- Minimize handling: The less touching, the better

- Dedicated transport: Don't mix sterile items with soiled/contaminated items

## Transport Methods

### Covered Carts

Most common method for transporting sterile items within the facility.

- Use dedicated carts for sterile items only

- Keep carts covered during transport

- Clean carts regularly

- Don't overload — items should not be compressed

### Case Carts

Enclosed carts prepared with all items needed for a specific surgical case.

- Provides protection during transport and storage

- Contents verified before leaving SPD

- Remains sealed until needed

### Transport Containers

For transporting items to off-site locations or within the facility.

- Must be clean and appropriate for sterile items

- Protects from environmental contamination

- Clearly labeled as sterile

⚠️ Never Transport Sterile and Soiled Together Sterile items must NEVER share transport with soiled or contaminated items. • Use separate carts • Use separate trips if necessary • If contamination occurs, the sterile item must be reprocessed

## Transport Routes

- Use designated clean corridors when possible

- Avoid passing through soiled areas

- Minimize time in transit

- Protect from environmental hazards (construction areas, outdoor exposure)

# 7.6 Case Cart Preparation

Case cart preparation ("picking") is assembling all the sterile items needed for a specific surgical procedure. Accuracy is essential.

## The Case Cart Process

### Receiving the Request

- Schedule or preference card specifies what is needed

- Pick list generated from the schedule

- Verify surgeon, procedure, date, time, location

### Gathering Items

- Review pick list: Understand what is needed before starting

- Locate items: Know your storage layout

- Verify each item: Correct item, correct quantity, intact package, valid date

- Inspect packages: Check integrity of every package

- Place in cart: Organize logically; protect packages

- Check off list: Mark each item as it is placed in cart

### Verification

- Double-check pick list against cart contents

- Verify quantities match

- Ensure all items are present

- Note any substitutions or shortages

🎯 Accuracy Prevents Delays Missing or incorrect items delay surgery and waste OR time. Take the time to verify each item carefully. If an item is unavailable: • Check alternate storage locations • Notify coordinator/charge • Document the issue • Provide approved substitute if available

## Case Cart Organization

- Heavy items on bottom: Instrument trays and containers

- Lighter items on top: Peel packs, small items

- Protect peel packs: Don't crush under heavy items

- Logical grouping: Items for same purpose together

- Pick list accessible: Attach pick list to cart

## Documentation

Complete documentation for each case cart:

- Date and time prepared

- Preparer identification (initials or ID)

- Case/patient information (per facility policy)

- Items included (pick list)

- Any substitutions or notes

# 7.7 Handling Recalls & Expired Items

Items may need to be removed from inventory due to recalls, expiration, or suspected problems. Proper handling protects patients.

## Recalls

A recall is notification that items must be removed from use due to a safety concern or defect.

### Types of Recalls

| SOURCE | EXAMPLES |
| --- | --- |
| Manufacturer recall | Defective product; packaging failure; contamination |
| FDA recall | Safety issue identified by regulatory authority |
| Internal recall | Positive BI; sterilizer malfunction; process failure identified |

### Recall Response

- Stop use immediately: Remove from service as soon as recall is identified

- Identify affected items: Use lot numbers, sterilization dates, or other identifiers

- Quarantine: Physically separate from usable inventory; clearly mark

- Notify appropriate parties: Supervisor, risk management, affected departments

- Document: Record all recall activities and affected items

- Follow up: Return to manufacturer, reprocess, or dispose per instructions

❤️ Traceability Enables Recalls This is why labeling and documentation are so important. If you can't identify which items came from which load, you cannot effectively recall. Complete, accurate records protect patients.

## Expired Items

Items that have exceeded their assigned shelf life (if facility uses date-based expiration).

### Handling Expired Items

- Remove from inventory immediately

- Do not use for patient care

- Evaluate packaging — if intact, may be reprocessed

- If packaging degraded, contents may need evaluation

- Document removal from inventory

### Preventing Expiration

- Implement FIFO consistently

- Monitor expiration dates during routine activities

- Adjust par levels if items consistently expire before use

- Consider shorter shelf life for slow-moving items

## Compromised Items

Items with suspected sterility compromise but no formal recall.

- When to remove: Any question about package integrity, storage conditions, or handling

- Principle: When in doubt, take it out

- Action: Remove, evaluate, reprocess or discard

- Documentation: Record reason for removal

# Knowledge Check: Module 7

Answer these questions to verify your understanding of this module. Discuss any questions with your preceptor.

### Section 7.1 — Storage Requirements

- What are the temperature and humidity limits for sterile storage areas?

Temperature: ________________  Humidity: ________________

- Why must the bottom shelf be at least 8-10 inches off the floor?

_____________________________________________________________________________

### Section 7.2 — Event-Related Sterility

- What does "event-related sterility" mean?

_____________________________________________________________________________

_____________________________________________________________________________

- List three events that compromise sterility:

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

### Section 7.3 — Package Inspection

- What should you check when inspecting packages before storage?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 7.4 — Stock Rotation

- What does FIFO stand for and how does it work?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 7.5 — Distribution

- What are three principles for transporting sterile items?

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

### Section 7.6 — Case Cart Preparation

- What should you verify for each item when picking a case cart?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 7.7 — Recalls & Expired Items

- What are the first three steps when a recall is identified?

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

# Skills Validation: Module 7

Your preceptor will observe you performing the following skills. Each skill must be demonstrated correctly.

| SKILL | MET | NOT MET | DATE | INITIALS |
| --- | --- | --- | --- | --- |
| Identifies sterile storage environmental requirements | ☐ | ☐ |  |  |
| Inspects packages for integrity before storage | ☐ | ☐ |  |  |
| Identifies and responds to wet pack correctly | ☐ | ☐ |  |  |
| Stores items correctly (location, orientation, spacing) | ☐ | ☐ |  |  |
| Applies FIFO principles when stocking shelves | ☐ | ☐ |  |  |
| Prepares case cart accurately from pick list | ☐ | ☐ |  |  |
| Verifies case cart contents before release | ☐ | ☐ |  |  |
| Transports sterile items appropriately | ☐ | ☐ |  |  |
| Identifies and removes expired items correctly | ☐ | ☐ |  |  |
| Responds appropriately to recall scenario | ☐ | ☐ |  |  |
| Documents storage and distribution activities | ☐ | ☐ |  |  |

### Module Completion

| Learner Name: | ________________________________________ |
| --- | --- |
| Learner Signature: | ________________________________________ |
| Date: | ________________________________________ |
| Preceptor Name: | ________________________________________ |
| Preceptor Signature: | ________________________________________ |
| Date: | ________________________________________ |

# Module 7 Summary

✓ Key Takeaways ✓ Sterile storage requires controlled environment (temp ≤75°F, humidity ≤70%) ✓ Shelving: bottom 8-10" off floor, top 18" from ceiling, 2" from walls ✓ Event-related sterility: items stay sterile until an event compromises the package ✓ Inspect every package before storage — external indicator, integrity, dryness, label ✓ Wet packs are NEVER acceptable — reprocess ✓ FIFO: First In, First Out — use oldest items first ✓ Protect sterile items during transport — covered, careful, clean ✓ Case cart accuracy prevents surgical delays ✓ Handle recalls immediately — stop, identify, quarantine, notify, document ✓ When in doubt about sterility, remove from service

## What's Next

In Module 8, you will learn about high-level disinfection and specialty processing — how to handle flexible endoscopes and other items that cannot be steam sterilized. You will learn:

- High-level disinfection principles

- Flexible endoscope reprocessing

- Automated endoscope reprocessors (AERs)

- Documentation and quality assurance for HLD

- Handling specialized equipment

════════════════════════════════════════

Storage protects the investment of sterilization.

Every package. Every inspection. Every time.

════════════════════════════════════════