SPD TECHNICIAN

TRAINING GUIDE

────────────────────────────────────────

MODULE 2

DECONTAMINATION

The Foundation of Safe Processing

IN THIS MODULE • PPE for Decontamination • Point-of-Use Treatment • Transport of Contaminated Items • Receiving & Sorting • Manual Cleaning Techniques • Ultrasonic Cleaning • Automated Washer Operation • Quality Verification

SIPS HEALTHCARE SOLUTIONS

Sterile By Design™ Curriculum

📚 LEARNING OBJECTIVES

After completing this module, you will be able to: 1. Don and doff decontamination PPE correctly 2. Explain the purpose and process of point-of-use treatment 3. Describe proper transport procedures for contaminated items 4. Perform manual cleaning using correct technique and sequence 5. Operate ultrasonic cleaning equipment safely and effectively 6. Load and operate automated washers according to procedure 7. Identify when items are clean and ready for inspection 8. Recognize common decontamination errors and how to avoid them

❤️ THE PATIENT CONNECTION Decontamination is where patient safety begins. If cleaning is inadequate, sterilization will fail. Every visible speck of blood or tissue you remove is a colony of bacteria that won't infect the next patient. This is not glamorous work. It is essential work.

# 2.1 PPE for Decontamination

The decontamination area is considered "dirty" — you will be handling items contaminated with blood, body fluids, and potentially infectious material. Proper PPE protects you from exposure.

## Required PPE

| PPE ITEM | PURPOSE | KEY POINTS |
| --- | --- | --- |
| Fluid-Resistant Gown | Protects clothing and skin from splashes | Must cover arms fully; ties in back; disposable or laundered |
| Heavy-Duty Gloves | Protects hands from sharps, chemicals, contamination | Thicker than exam gloves; check for tears before use; change when torn |
| Face Shield OR Goggles + Mask | Protects eyes, nose, mouth from splashes | Full face shield preferred; if goggles, must add surgical mask |
| Fluid-Resistant Shoe Covers | Protects shoes; prevents tracking contamination | Cover entire shoe; change if wet or visibly soiled |
| Hair Cover (Bouffant) | Contains hair; prevents contamination | Must cover all hair; replace if it falls off |
| Sleeve Protectors (as needed) | Additional arm protection for heavy splash work | Used in addition to gown when extended arm immersion expected |

## Donning Sequence (Putting On)

Put on PPE in this order to maximize protection:

| STEP 1 | Hair Cover Place bouffant cap over all hair. Tuck in any stray hairs. |
| --- | --- |

| STEP 2 | Shoe Covers Place over shoes, ensuring complete coverage. Stand on clean surface after applying. |
| --- | --- |

| STEP 3 | Gown Put on fluid-resistant gown. Tie at neck and waist. Ensure arms are fully covered. |
| --- | --- |

| STEP 4 | Mask/Respirator (if separate) Place over nose and mouth. Secure ties or ear loops. Mold nose piece. |
| --- | --- |

| STEP 5 | Eye Protection Put on goggles or face shield. Ensure seal around eyes. |
| --- | --- |

| STEP 6 | Gloves Put on heavy-duty gloves LAST. Pull glove cuffs over gown cuffs. |
| --- | --- |
| 💡 Gloves Over Gown Cuffs Always pull your glove cuffs OVER your gown cuffs. This prevents fluids from running down into your gloves and onto your skin. |

## Doffing Sequence (Taking Off)

Remove PPE in this order to prevent self-contamination. The outside of your PPE is contaminated — avoid touching it with bare hands.

| STEP 1 | Gloves • Grasp outside of one glove at wrist • Peel off, turning inside out • Hold removed glove in gloved hand • Slide finger under remaining glove at wrist • Peel off over first glove, creating a ball • Dispose in appropriate container |
| --- | --- |

→ PERFORM HAND HYGIENE

| STEP 2 | Face Shield/Goggles • Grasp the strap or earpieces (not the front) • Lift away from face • Dispose or place in designated bin for cleaning |
| --- | --- |

| STEP 3 | Gown • Unfasten ties • Pull gown away from body, touching only inside • Turn inside out while removing • Roll into a bundle • Dispose in appropriate container |
| --- | --- |

| STEP 4 | Mask (if separate) • Grasp ties or ear loops only (not front of mask) • Remove by lifting away from face • Dispose |
| --- | --- |

| STEP 5 | Shoe Covers • Remove one at a time • Step onto clean surface with uncovered foot • Remove second cover • Dispose |
| --- | --- |

| STEP 6 | Hair Cover • Grasp top of bouffant • Lift off head • Dispose |
| --- | --- |

→ PERFORM HAND HYGIENE AGAIN

⚠️ The Golden Rule of Doffing The OUTSIDE of your PPE is contaminated. The INSIDE is clean. Always touch only the inside or the ties/straps when removing.

# 2.2 Point-of-Use Treatment

Point-of-use treatment refers to the actions taken at the location where instruments were used (usually the OR) before they are transported to SPD. This is not your primary responsibility — it happens in the OR — but you need to understand it because it directly affects your work.

## Why Point-of-Use Treatment Matters

Blood and tissue begin to dry on instruments within minutes. Once dried, this organic matter becomes much harder to remove. Bioburden that is allowed to dry:

- Adheres more strongly to instrument surfaces

- Can form biofilm that protects bacteria

- May not be fully removed by standard cleaning

- Can compromise sterilization

📌 The Drying Problem Studies show that instruments are significantly harder to clean if left dry for more than 15-20 minutes. Point-of-use treatment buys time by keeping instruments moist until they reach decontamination.

## Point-of-Use Treatment Methods

| METHOD | DESCRIPTION | KEY POINTS |
| --- | --- | --- |
| Pre-Treatment Spray | Enzymatic foam or spray applied to instruments immediately after use | Keeps instruments moist; begins breaking down organic matter; follow manufacturer contact time |
| Pre-Treatment Gel | Thick gel that maintains moisture on instruments | Good for longer transport times; covers lumens and box locks |
| Moist Towels | Damp towels placed over instruments in transport container | Simple but less effective than enzymatic products |
| Soaking (Limited Use) | Instruments placed in solution | Generally discouraged — can damage instruments, staff safety concerns with transport |

## What to Look for When Receiving Items

When items arrive in decontam, assess whether point-of-use treatment was adequate:

- Dried blood/tissue: Heavy dried matter indicates delayed transport or no pre-treatment — will require extra attention

- Moist instruments: Good sign — pre-treatment or quick transport

- Pre-treatment residue: Enzymatic foam or gel visible — treatment was applied

- Gross soil remaining: Large pieces of tissue or heavy contamination — may need extended soaking

🎯 SBD Communication If items consistently arrive with heavy dried soil, this is a system problem that should be escalated: "Issue identified: Instruments from [location] arriving with excessive dried bioburden. Escalating to supervisor."

# 2.3 Transport of Contaminated Items

Contaminated instruments must be transported safely from the point of use to the decontamination area. Transport protects staff, prevents environmental contamination, and maintains conditions that support effective cleaning.

## Transport Container Requirements

- Leak-proof: Must contain any fluids without leaking

- Puncture-resistant: Must protect handlers from sharps

- Covered/Enclosed: Must prevent exposure of contents during transport

- Labeled: Must be clearly marked as biohazardous

- Cleanable: Must be able to be decontaminated after each use

## Transport Principles

| DO | DON'T |
| --- | --- |
| Use designated transport carts and containers | Carry loose instruments by hand |
| Keep containers closed during transport | Leave containers open in hallways |
| Use designated routes (contaminated corridors) | Transport through clean areas or patient areas |
| Transport promptly to minimize drying time | Let carts sit for extended periods |
| Separate sharps or secure them properly | Pile instruments loosely where sharps can shift |
| Clean transport containers after each use | Reuse containers without cleaning |

## Receiving Contaminated Items

When items arrive in the decontamination area:

| STEP 1 | Verify PPE Ensure you are fully donned in appropriate decontam PPE before handling any items. |
| --- | --- |

| STEP 2 | Receive at Designated Area Accept items only at the designated receiving point, never in clean areas. |
| --- | --- |

| STEP 3 | Inspect Container Check for leaks, damage, or improperly closed containers. Report issues. |
| --- | --- |

| STEP 4 | Document Receipt • Scan or log incoming items per facility protocol • Note time received • Note any issues with transport condition |
| --- | --- |

| STEP 5 | Transfer to Decontam Workflow Move items to the appropriate processing station based on type and priority. |
| --- | --- |

# 2.4 Receiving & Sorting

Before cleaning can begin, items must be properly sorted. Sorting ensures that each item receives appropriate processing and helps organize your workflow.

## Sorting Criteria

### By Contamination Level

- Heavy soil: Gross contamination, dried blood/tissue — may require pre-soaking

- Moderate soil: Visible contamination but not excessive

- Light soil: Minimal visible contamination

### By Instrument Type

- General instruments: Standard surgical instruments

- Lumened instruments: Cannulas, suction tips, scope channels — require special attention

- Power equipment: Drills, saws — may have specific cleaning requirements

- Delicate instruments: Microsurgical, ophthalmic — require gentle handling

- Complex devices: Items requiring disassembly

### By Processing Method

- Manual cleaning only: Items that cannot go through automated washers

- Automated washer compatible: Most general instruments

- Special processing: Items with specific IFU requirements

## Critical Sorting Steps

| STEP 1 | Open All Hinged Instruments Box locks, scissors, clamps must be open to allow cleaning of all surfaces. |
| --- | --- |

| STEP 2 | Disassemble Multi-Part Instruments Follow IFU for proper disassembly. Keep parts together. |
| --- | --- |

| STEP 3 | Separate Sharps Identify and segregate scalpels, needles, and pointed instruments for safe handling. |
| --- | --- |

| STEP 4 | Identify Items Needing Special Processing Check for devices with specific IFU requirements. |
| --- | --- |

| STEP 5 | Remove Disposable Items Discard single-use items per facility protocol. |
| --- | --- |
| 🛡️ Sharp Safety During Sorting LOOK before you reach into any container Use forceps or a magnetic retrieval tool to handle sharps Never reach into containers you cannot see into clearly Keep sharps pointed away from yourself and others |

# 2.5 Manual Cleaning

Manual cleaning is the hands-on removal of soil from instruments using water, detergent, and mechanical action (brushing, wiping). Even items that will go through automated washers often require some manual pre-cleaning.

❤️ The Most Important Step Manual cleaning is often called the most important step in reprocessing. No disinfection or sterilization method can overcome inadequate cleaning. Take your time. Do it right.

## Manual Cleaning Supplies

| ITEM | PURPOSE | KEY POINTS |
| --- | --- | --- |
| Enzymatic Detergent | Breaks down blood, protein, fat, carbohydrates | Mix per manufacturer instructions; use correct temperature; respect contact time |
| Brushes (Various Sizes) | Mechanical removal of soil from surfaces | Match brush size to instrument; use for serrations, box locks, lumens; disposable or sterilizable |
| Lumen Brushes/Stylets | Clean inside channels and lumens | Must be appropriate diameter; pass completely through |
| Syringes for Flushing | Force solution through lumens | Use appropriate volume; must see clear solution exit |
| Low-Lint Cloths | Wiping surfaces | Use for wiping, not for scrubbing; discard when soiled |
| Utility Sink | Rinsing and cleaning work surface | Maintain water quality; clean sink regularly |

## Manual Cleaning Procedure

| STEP 1 | Prepare Cleaning Solution • Fill sink or basin with water at correct temperature (per detergent IFU) • Add enzymatic detergent at correct dilution • Verify temperature with thermometer • Mix thoroughly |
| --- | --- |

| STEP 2 | Pre-Rinse (Gross Soil Removal) • Hold instrument under cold running water • Remove large debris • Flush lumens with water • Do NOT use hot water — it can set proteins |
| --- | --- |

| STEP 3 | Submerge in Enzymatic Solution • Place instruments in solution with hinges open • Ensure complete submersion • Respect minimum contact time per detergent IFU • Agitate periodically to ensure solution reaches all surfaces |
| --- | --- |

| STEP 4 | Brush All Surfaces • Work under the water surface to prevent aerosolization • Brush all surfaces — jaws, serrations, box locks, ratchets • Use appropriately sized brushes for each area • Pay special attention to crevices and hard-to-reach areas |
| --- | --- |

| STEP 5 | Clean Lumens • Pass appropriately sized brush through entire lumen • Flush with enzymatic solution using syringe • Repeat until visual inspection shows no soil • Final flush must show clear solution |
| --- | --- |

| STEP 6 | Rinse Thoroughly • Rinse all surfaces with clean water • Flush all lumens with clean water • Remove all detergent residue • Use water quality specified by facility (often DI or treated water) |
| --- | --- |

| STEP 7 | Inspect • Visually inspect each instrument • Use magnification for difficult areas • If soil remains, repeat cleaning • If clean, proceed to next step (automated washing or drying) |
| --- | --- |
| ⚠️ Critical Points • Always work UNDER the water surface to prevent splashing and aerosols • Never scrub so hard you damage the instrument • Change solutions when visibly soiled or per facility policy • If you can't get it clean, don't send it forward — escalate |

## Cleaning Specific Areas

### Serrations

Serrations are the grooved gripping surfaces on forceps and clamps. Soil can hide between the ridges.

- Use a brush with bristles that can get between serration ridges

- Brush in direction of serrations, not across them

- Inspect with magnification — bioburden hides here

### Box Locks

Box locks are the hinged areas where instrument jaws pivot. Blood and debris accumulate in the crevices.

- Open instrument fully during cleaning

- Use a narrow brush to clean inside the box lock

- Work the instrument open and closed while submerged to release trapped debris

### Ratchets

Ratchets are the locking teeth that hold instruments closed. Debris collects between teeth.

- Brush along the teeth, not against them

- Work with instrument in different ratchet positions

- Inspect each tooth surface

### Lumens

Lumens are the internal channels of hollow instruments. These are the most challenging areas to clean.

- Brush must pass completely through — not just inserted and withdrawn

- Brush diameter must be appropriate for lumen size (follow IFU)

- Flush with syringe until effluent is clear

- Some lumens require pressurized flushing

### Cannulated Instruments

Cannulated instruments (drills, screwdrivers) have channels that are often overlooked.

- Follow manufacturer IFU for each device

- Use specified cleaning accessories

- Disassemble as directed

- Some require specialized brushes or adapters for flushing

# 2.6 Ultrasonic Cleaning

Ultrasonic cleaning uses high-frequency sound waves to create microscopic bubbles in a cleaning solution. These bubbles implode (a process called cavitation), creating tiny shock waves that dislodge soil from instrument surfaces — including areas that brushes cannot reach.

## When to Use Ultrasonic Cleaning

- Instruments with complex surfaces (serrations, box locks, ratchets)

- Instruments with blind holes or narrow lumens that brushes can't access

- Delicate instruments where aggressive brushing could cause damage

- After manual cleaning to enhance effectiveness

- When specified by manufacturer IFU

⚠️ Not a Substitute for Manual Cleaning Ultrasonic cleaning enhances manual cleaning — it does not replace it. Heavy soil should be removed manually before ultrasonic processing.

## How Ultrasonic Cleaning Works

### Cavitation

The ultrasonic transducers create sound waves that pass through the liquid. These waves create alternating high-pressure and low-pressure cycles. During low-pressure cycles, microscopic bubbles form. During high-pressure cycles, these bubbles collapse violently, creating intense localized cleaning action.

### What Affects Effectiveness

| FACTOR | IMPACT | BEST PRACTICE |
| --- | --- | --- |
| Temperature | Affects cavitation efficiency and enzyme activity | Follow detergent IFU; typically 100-130°F (40-55°C) |
| Detergent | Reduces surface tension, enhances cleaning | Use manufacturer-recommended enzymatic detergent; correct dilution |
| Contact Time | Longer time = more cleaning action | Follow IFU; typically 3-10 minutes |
| Loading | Overloading reduces effectiveness | Don't overcrowd; ensure solution contact with all surfaces |
| Water Quality | Minerals can affect cavitation | Use treated water per facility protocol |

## Ultrasonic Cleaning Procedure

| STEP 1 | Prepare the Ultrasonic Cleaner • Fill with water to appropriate level • Add enzymatic detergent at correct dilution • Verify temperature is in acceptable range • Run degassing cycle (removes dissolved air that interferes with cavitation) |
| --- | --- |

| STEP 2 | Prepare Instruments • Complete manual pre-cleaning first • Open all hinged instruments • Disassemble multi-part items • Place in ultrasonic basket or tray |
| --- | --- |

| STEP 3 | Load the Unit • Submerge instruments completely • Do not overload — leave space between items • Keep similar metals together (prevent galvanic corrosion) • Heavy items on bottom; delicate items on top |
| --- | --- |

| STEP 4 | Run the Cycle • Set timer per manufacturer IFU • Cover the tank (reduces aerosols and maintains temperature) • Do not add or remove items during cycle • Allow full cycle to complete |
| --- | --- |

| STEP 5 | Post-Cycle • Remove instruments • Rinse thoroughly to remove detergent residue • Inspect for remaining soil • If soil remains, repeat or escalate • Proceed to automated washing or next step |
| --- | --- |

## Ultrasonic Testing

Ultrasonic cleaners must be tested regularly to verify they are producing adequate cavitation.

### Aluminum Foil Test

A simple in-house test using aluminum foil:

- Suspend a sheet of aluminum foil vertically in the tank

- Run the ultrasonic cleaner for 20-30 seconds

- Examine the foil for pitting and holes

- Adequate cavitation = even distribution of pitting across the foil

### Commercial Cavitation Test Devices

Commercial test strips or devices may also be used per manufacturer instructions. Document all testing per facility protocol.

💡 Maintenance • Change solution per manufacturer/facility guidelines (typically each use or every 2-4 hours of use) • Clean and drain tank at end of day • Test cavitation at specified intervals • Report any unusual sounds, leaks, or performance issues

# 2.7 Automated Washer Operation

Automated washers (also called washer-disinfectors or washer-decontaminators) provide standardized, reproducible cleaning and thermal disinfection of instruments. They are more consistent than manual cleaning alone and provide documented cycle parameters.

## How Automated Washers Work

Automated washers move instruments through a series of phases:

| PHASE | PURPOSE | DETAILS |
| --- | --- | --- |
| 1. Pre-Wash (Cold) | Remove gross soil without setting proteins | Cold water (~70°F / 21°C); no detergent |
| 2. Enzyme Wash (Warm) | Enzymatic breakdown of organic soil | Warm water (per detergent IFU); enzymatic detergent added; agitation |
| 3. Detergent Wash (Hot) | Further cleaning with surfactant detergent | Hot water; detergent; mechanical agitation |
| 4. Rinse (Multiple) | Remove detergent residue | Multiple rinse phases; treated water; progressively hotter |
| 5. Thermal Disinfection | High-level disinfection through heat | Hot water (180°F / 82°C minimum) for specified time; typically 1-10 minutes |
| 6. Drying | Remove moisture | Hot air circulation; some units use vacuum |

## Loading the Automated Washer

Proper loading is critical to cleaning effectiveness.

### General Loading Principles

- Open hinged instruments: All box locks, ratchets, and jaws must be open

- Disassemble: Multi-part instruments must be taken apart

- Position for drainage: Items should drain freely — no cupping

- Don't overcrowd: Water and air must reach all surfaces

- Secure items: Nothing should come loose and block spray arms

- Heavy items on bottom: Light items can blow around

### Special Considerations

- Lumened instruments: Must be connected to flushing ports or positioned for water entry

- Containers/Basins: Place inverted or at angle to prevent pooling

- Delicate items: Use appropriate racks; may need gentler cycle

💡 Verify Spray Arm Clearance Before closing the washer, manually rotate the spray arms to ensure nothing will obstruct them during the cycle.

## Operating the Washer

| STEP 1 | Pre-Cycle Checks • Verify detergent and rinse agent levels • Check that spray arms rotate freely • Ensure drain is clear • Verify no items from previous load remain |
| --- | --- |

| STEP 2 | Load Instruments • Place items in racks per loading guidelines • Connect lumened instruments to flush ports • Verify nothing blocks spray arms • Close door securely |
| --- | --- |

| STEP 3 | Select Cycle • Choose appropriate cycle for items being processed • General, delicate, or specialty cycles as appropriate • Never skip phases or shorten cycles |
| --- | --- |

| STEP 4 | Initiate Cycle • Start the cycle • Document load contents per facility protocol • Note cycle start time |
| --- | --- |

| STEP 5 | Monitor Cycle • Most units run unattended • Respond promptly to alarms • Do not open door during cycle unless emergency |
| --- | --- |

| STEP 6 | Cycle Completion • Allow cycle to complete fully • Review printout or display for cycle parameters • Verify thermal disinfection parameters were met • If parameters not met, do not use items — troubleshoot and rewash |
| --- | --- |

| STEP 7 | Unload • Open door carefully (hot items, steam) • Remove racks • Inspect items for cleanliness and dryness • Items still wet should be dried before packaging |
| --- | --- |

## Thermal Disinfection Parameters

Thermal disinfection occurs when instruments are exposed to hot water for a sufficient time. This is measured using the A₀ concept:

📌 Understanding A₀ A₀ (A-not or A-zero) is a measure of thermal disinfection effectiveness. A₀ 600 is the minimum recommended for general medical devices. A₀ 3000 may be required for items with higher bioburden or infection risk. Your washer calculates and displays A₀ automatically.

Example thermal disinfection parameters:

| TEMPERATURE | TIME | A₀ ACHIEVED |
| --- | --- | --- |
| 180°F (82°C) | 10 minutes | A₀ 600 |
| 194°F (90°C) | 1 minute | A₀ 600 |
| 194°F (90°C) | 5 minutes | A₀ 3000 |

## Common Washer Problems

| PROBLEM | POSSIBLE CAUSES | ACTION |
| --- | --- | --- |
| Items not clean | Overloading; improper loading; detergent issue; cycle problem | Re-wash; check loading; verify detergent; escalate if persistent |
| Items still wet | Cupping (water pooling); drying phase inadequate | Reposition for drainage; verify drying cycle; hand dry if needed |
| White residue on items | Hard water deposits; detergent residue | Check water treatment; verify rinse agent; check detergent dilution |
| Thermal parameters not met | Water heater issue; sensor failure; loading issue | Do not use items; rewash after issue resolved; escalate to biomed |
| Cycle aborts | Door not latched; system error; overload | Check door; reduce load; restart; escalate if continues |

# 2.8 Quality Verification

Decontamination is only effective if cleaning is verified. Multiple verification methods are used to ensure items are truly clean before moving to the next phase.

## Visual Inspection

Every instrument should be visually inspected after cleaning. This is the most important verification step.

### What to Look For

- Visible soil: Blood, tissue, debris — if you see it, it's not clean

- Staining: May indicate residue or corrosion

- Debris in crevices: Check box locks, serrations, ratchets, lumens

- Residue on surfaces: Film or residue indicates incomplete rinsing

### Use Magnification

The naked eye is inadequate for thorough inspection. Use:

- Lighted magnification lamp (2x-5x)

- Borescope or endoscope camera for lumens

- Adequate lighting (direct, shadowless)

🎯 SBD Standard If you cannot verify that an item is clean, it is not clean. Never send an item forward if you are uncertain about its cleanliness. When in doubt, re-clean.

## Protein Detection Testing

Protein detection tests provide objective verification that organic soil has been removed. They detect protein residue that may not be visible to the eye.

### How It Works

- Test swabs or devices contain chemicals that react with protein

- A color change indicates protein is present

- Used on surfaces, crevices, and lumens

### When to Use

- Random sampling per facility protocol

- When visual inspection is difficult (complex devices)

- After cleaning process changes

- To validate cleaning of high-risk items

### Interpreting Results

- Negative (no color change): Protein not detected — item may proceed

- Positive (color change): Protein detected — re-clean and retest

## ATP Testing

ATP (adenosine triphosphate) testing detects the presence of organic material by measuring ATP, which is present in all living cells and organic matter.

### How It Works

- Swab is used to sample a surface

- Swab is placed in a luminometer device

- Device measures ATP and displays a numerical result (RLU — relative light units)

### Interpreting Results

- Results below threshold = acceptable cleaning

- Results above threshold = inadequate cleaning — re-clean and retest

- Thresholds are established by facility based on device validation

## Washer Testing

Automated washers must be tested regularly to verify they are functioning correctly.

| TEST | PURPOSE | FREQUENCY |
| --- | --- | --- |
| Temperature verification | Confirms thermal disinfection parameters are met | Each load (via printout); external verification per PM schedule |
| Cleaning efficacy test | Verifies mechanical cleaning is effective | Per manufacturer recommendation; typically daily or weekly |
| Detergent concentration | Verifies proper chemical dilution | Per facility protocol; often daily |
| Water quality testing | Verifies water treatment is adequate | Per facility protocol |

## Documentation

All verification activities must be documented:

- Washer cycle printouts (maintained per facility retention policy)

- Washer testing logs

- Protein/ATP testing results

- Any items that failed inspection and required re-processing

- Any equipment issues or unusual findings

🎯 If It Isn't Documented, It Didn't Happen This is a core SBD principle. Your documentation provides evidence that you followed proper procedures. Document in real time — not at the end of your shift.

# 2.9 Common Decontamination Errors

Learning from common errors helps you avoid them. These are mistakes that experienced technicians still make — stay vigilant.

| ERROR | WHY IT MATTERS | PREVENTION |
| --- | --- | --- |
| Not opening hinged instruments | Box locks trap blood and tissue; cannot be cleaned closed | ALWAYS open every hinged instrument before processing |
| Skipping manual pre-cleaning | Automated washers cannot remove heavy or dried soil | Pre-clean all items; never put heavily soiled items directly in washer |
| Using hot water for pre-rinse | Hot water sets proteins, making them harder to remove | Use COLD water for initial gross soil removal |
| Incorrect detergent dilution | Too little = inadequate cleaning; too much = residue, damage | Follow manufacturer instructions precisely; measure, don't guess |
| Overloading washers | Water and air cannot reach all surfaces | Leave space between items; follow manufacturer loading guidelines |
| Not flushing lumens | Bioburden remains inside channels | Brush AND flush all lumens; verify clear effluent |
| Brushing above water surface | Creates aerosols containing potentially infectious material | ALWAYS brush submerged |
| Inadequate inspection | Soil passes to assembly/sterilization undetected | Use magnification; inspect every item; when in doubt, re-clean |
| Mixing dissimilar metals | Can cause galvanic corrosion and damage | Separate stainless steel, aluminum, plated instruments |
| Ignoring IFUs | Manufacturer-specific requirements not met | Check IFU for unfamiliar items; follow specific requirements |

⚠️ The Consequence of Decontamination Errors When decontamination fails, sterilization cannot fully compensate. Bioburden left on instruments can survive the sterilization process. Patients can develop life-threatening infections from improperly processed instruments. Every error in decontam is a potential patient safety event.

# Knowledge Check: Module 2

Answer these questions to verify your understanding of this module. Discuss any questions with your preceptor.

### Section 2.1 — PPE

- List the required PPE for the decontamination area:

_____________________________________________________________________________

_____________________________________________________________________________

- What is the correct sequence for doffing (removing) PPE?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 2.2 & 2.3 — Point-of-Use & Transport

- Why is point-of-use treatment important?

_____________________________________________________________________________

_____________________________________________________________________________

- What are the requirements for transport containers?

_____________________________________________________________________________

### Section 2.5 — Manual Cleaning

- Why should cold water (not hot) be used for the initial pre-rinse?

_____________________________________________________________________________

- Why must you brush instruments UNDER the water surface?

_____________________________________________________________________________

- List three areas of instruments that require special attention during manual cleaning:

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

### Section 2.6 & 2.7 — Ultrasonic & Automated Washers

- What is cavitation and how does it clean instruments?

_____________________________________________________________________________

_____________________________________________________________________________

- What are the key principles for loading an automated washer?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 2.8 — Quality Verification

- What should you do if you find an item that may not be clean after automated washing?

_____________________________________________________________________________

- What is the purpose of protein detection testing?

_____________________________________________________________________________

### Section 2.9 — Common Errors

- List three common decontamination errors and their consequences:

1. Error: _________________________  Consequence: _____________________________

2. Error: _________________________  Consequence: _____________________________

3. Error: _________________________  Consequence: _____________________________

# Skills Validation: Module 2

Your preceptor will observe you performing the following skills. Each skill must be demonstrated correctly.

| SKILL | MET | NOT MET | DATE | INITIALS |
| --- | --- | --- | --- | --- |
| Correctly dons PPE for decontam in proper sequence | ☐ | ☐ |  |  |
| Correctly doffs PPE without self-contamination | ☐ | ☐ |  |  |
| Properly receives and sorts contaminated items | ☐ | ☐ |  |  |
| Prepares enzymatic cleaning solution correctly | ☐ | ☐ |  |  |
| Performs manual cleaning technique correctly | ☐ | ☐ |  |  |
| Properly cleans lumened instruments | ☐ | ☐ |  |  |
| Operates ultrasonic cleaner correctly | ☐ | ☐ |  |  |
| Loads automated washer correctly | ☐ | ☐ |  |  |
| Operates automated washer and monitors cycle | ☐ | ☐ |  |  |
| Inspects items for cleanliness using magnification | ☐ | ☐ |  |  |
| Documents decontamination activities correctly | ☐ | ☐ |  |  |
| Demonstrates safe sharps handling | ☐ | ☐ |  |  |

### Module Completion

| Learner Name: | ________________________________________ |
| --- | --- |
| Learner Signature: | ________________________________________ |
| Date: | ________________________________________ |
| Preceptor Name: | ________________________________________ |
| Preceptor Signature: | ________________________________________ |
| Date: | ________________________________________ |

# Module 2 Summary

✓ Key Takeaways ✓ PPE protects you — don and doff correctly every time ✓ Point-of-use treatment prevents soil from drying and becoming harder to clean ✓ Manual cleaning is the most important step — there are no shortcuts ✓ Always brush UNDER the water surface to prevent aerosols ✓ Cold water for pre-rinse (hot water sets proteins) ✓ Ultrasonic cleaning enhances but does not replace manual cleaning ✓ Proper loading of automated washers is critical to effectiveness ✓ Every item must be inspected — if you're unsure, re-clean ✓ Document everything in real time

## What's Next

In Module 3, you will learn how to inspect instruments — a critical step between cleaning and assembly. You will learn:

- Visual inspection techniques

- Functional testing of instruments

- Identifying damage and wear

- Instrument identification methods

- When to remove items from service

════════════════════════════════════════

Cleaning is where safety begins.

If it isn't clean, it cannot be sterile.

════════════════════════════════════════