SPD TECHNICIAN

TRAINING GUIDE

────────────────────────────────────────

MODULE 3

INSPECTION &

IDENTIFICATION

The Critical Eye of Quality

IN THIS MODULE • The Purpose of Inspection • Visual Inspection Techniques • Functional Testing • Identifying Damage & Defects • Instrument Identification Methods • Basic Instrument Categories • Removing Items from Service • Lubrication & Instrument Care

SIPS HEALTHCARE SOLUTIONS

Sterile By Design™ Curriculum

📚 LEARNING OBJECTIVES

After completing this module, you will be able to: 1. Explain why inspection is critical to patient safety 2. Perform visual inspection using proper technique and magnification 3. Conduct functional testing of common instrument types 4. Identify common types of instrument damage and defects 5. Use multiple methods to identify instruments (markings, catalogs, resources) 6. Categorize instruments by function (cutting, clamping, retracting, etc.) 7. Determine when an instrument must be removed from service 8. Apply instrument lubricant correctly

❤️ THE PATIENT CONNECTION Inspection is your last chance to catch problems before an instrument goes to a patient. A damaged instrument can fail during surgery. A dirty instrument can cause infection. Your eyes — and your willingness to reject defective items — protect patients.

# 3.1 The Purpose of Inspection

Inspection is the bridge between cleaning and assembly. It serves three critical functions:

### 1. Verify Cleanliness

Confirm that decontamination was effective. No visible soil, staining, or debris should remain. If an item isn't clean, it cannot be sterilized effectively.

### 2. Verify Function

Confirm that the instrument works properly. Damaged or malfunctioning instruments can fail during surgery, putting patients at risk.

### 3. Verify Condition

Confirm that the instrument is in acceptable condition. Corrosion, pitting, cracks, and other defects can harbor bacteria or cause instrument failure.

🎯 The Three Questions of Inspection Ask yourself three questions about every instrument you inspect: 1. IS IT CLEAN? (No visible soil, staining, or debris) 2. DOES IT WORK? (Functions as intended) 3. IS IT SAFE? (No damage that could harm patient or compromise sterility) If the answer to ANY question is "no" or "I'm not sure" — the instrument does not pass inspection.

## When Inspection Occurs

Inspection happens at multiple points in the workflow:

| INSPECTION POINT | PURPOSE | ACTION IF PROBLEM FOUND |
| --- | --- | --- |
| After Decontamination | Verify cleaning was effective; initial function check | Re-clean if soil remains; set aside if damaged |
| During Assembly | Verify each instrument before placing in tray | Do not include in tray; replace with functional item |
| After Sterilization | Verify packaging integrity; sterility indicators | Do not use; reprocess or discard |
| At Point of Use (OR) | Final verification before patient use | Remove from field; request replacement |

Your primary inspection responsibility is after decontamination and during assembly. However, understand that inspection is a continuous process throughout the instrument lifecycle.

# 3.2 Visual Inspection Techniques

Visual inspection requires more than just a quick glance. Proper technique ensures you catch problems that could be missed with casual observation.

## The Inspection Environment

### Lighting

Adequate lighting is essential. You cannot see what you cannot illuminate.

- Direct, bright light: Use a dedicated inspection light or lighted magnifier

- Shadow-free: Position light to eliminate shadows in crevices

- Consistent: Maintain the same lighting conditions for all inspection

### Magnification

The naked eye is insufficient for thorough inspection. Always use magnification.

- Lighted magnifier: 2x-5x magnification with integrated light

- Magnifying lamp: Hands-free option for detailed work

- Borescope: For inspecting lumens and internal channels

⚠️ Magnification Is Not Optional Studies show that visual inspection without magnification misses a significant percentage of residual soil. What looks clean to the naked eye may harbor contamination visible only under magnification. Use magnification for EVERY inspection.

## Systematic Inspection Process

Use a systematic approach to ensure you examine every part of each instrument.

### The SIPS Inspection Sequence

- Overall condition: Look at the entire instrument — obvious damage, corrosion, discoloration

- Working surfaces: Examine the functional areas — jaws, tips, cutting edges, serrations

- Joints and hinges: Inspect box locks, pivot points, ratchets

- Handles and shafts: Check for cracks, bends, loose parts

- Lumens and channels: Use borescope or flush test for hollow instruments

- Markings and labels: Verify legibility of identification markings

## What to Look For: Cleanliness

| FINDING | DESCRIPTION | ACTION |
| --- | --- | --- |
| Visible soil | Blood, tissue, debris visible on any surface | Return to decontamination; do not proceed |
| Staining | Discoloration that may indicate residue | Re-clean; investigate cause if persistent |
| Debris in crevices | Material trapped in box locks, serrations, ratchets | Return to decontamination; may need manual attention |
| Lumen obstruction | Channel not clear or flush test fails | Return to decontamination; brush and flush lumen |
| Water spots | Mineral deposits from water (usually cosmetic) | Wipe with cloth; check water quality if persistent |
| Detergent residue | Film or white residue from cleaning chemicals | Re-rinse; check rinse cycle; check detergent dilution |

# 3.3 Functional Testing

Functional testing verifies that an instrument works as designed. A clean instrument that doesn't function properly is still unusable.

## Testing Hinged Instruments

Hinged instruments (scissors, hemostats, needle holders, clamps) have moving parts that must work smoothly.

### Box Lock Test

- Open and close fully: Should move smoothly throughout entire range of motion

- No stiffness: Excessive resistance indicates damage, debris, or need for lubrication

- No looseness: Excessive play indicates worn box lock

- No grinding: Grinding sensation indicates damage or debris

### Ratchet Test

- Engage each position: Close instrument to each ratchet position

- Secure hold: Instrument should stay closed at each position without slipping

- Positive click: Should feel and hear distinct engagement

- Release test: Should release without excessive force

### Jaw Alignment Test

- Close jaws: Tips should meet evenly

- No gaps: Jaws should close completely with no light visible between them

- Parallel closure: Jaws should meet along their full length

## Testing Scissors

Scissors have specific testing requirements beyond general hinged instruments.

### Cutting Test

| SCISSOR TYPE | TEST MATERIAL | EXPECTED RESULT |
| --- | --- | --- |
| General surgical scissors | 4 layers of gauze or test material | Clean cut to the tip; no snagging or tearing |
| Delicate/fine scissors | Single layer of tissue paper or 2 layers gauze | Clean cut; no fraying at the tip |
| Heavy scissors (Mayo) | Multiple layers of gauze | Clean cut with minimal effort |
| Suture scissors | Suture material (per facility protocol) | Clean cut of suture without snagging |
| 💡 Blade Inspection Beyond the cutting test, visually inspect scissor blades for: • Nicks or chips in the cutting edge • Burred edges (rolled or rough edges) • Blade separation (tips don't meet) • Corrosion on cutting surfaces |

## Testing Needle Holders

Needle holders must grip needles securely without damaging them.

### Grip Test

- Insert test pin or needle: Use appropriate size test device

- Close to first ratchet: Needle should be held securely

- Rotate test: Needle should not rotate within jaws

- Pull test: Needle should not pull out easily

### Jaw Surface Inspection

- Tungsten carbide inserts: Check for wear, chips, missing sections

- Cross-hatching: Pattern should be intact and distinct

- Jaw alignment: Surfaces should meet evenly

## Testing Other Instruments

### Forceps (Thumb Forceps)

- Spring tension: Should spring open when released

- Tip alignment: Tips should meet precisely

- Teeth engagement: Teeth should interlock properly (if toothed)

### Retractors

- Blade condition: No bends, cracks, or sharp burrs

- Ratchet function: Self-retaining retractors hold position

- Screw mechanisms: Turn freely; hold adjustment

### Rongeurs and Punches

- Bite test: Jaws should close completely

- Cutting edges: Sharp, no nicks or chips

- Spring return: Handles return to open position

# 3.4 Identifying Damage & Defects

Knowing what damage looks like is essential to identifying it. This section covers common types of instrument damage and defects.

## Types of Damage

### Corrosion

| TYPE | APPEARANCE | COMMON CAUSES |
| --- | --- | --- |
| Rust | Orange-brown discoloration; rough texture | Moisture exposure; improper drying; damaged surface |
| Pitting | Small holes or craters in surface | Chemical attack; prolonged moisture; saline exposure |
| Staining | Discoloration (brown, black, rainbow) | Chemical residue; mineral deposits; heat damage |
| Galvanic corrosion | Corrosion at junction of different metals | Mixing dissimilar metals during processing |

### Physical Damage

| TYPE | APPEARANCE | COMMON CAUSES |
| --- | --- | --- |
| Cracks | Visible fractures in metal; may be hairline | Stress; improper use; manufacturing defect |
| Bent/misaligned | Instrument not straight; jaws don't meet | Improper use; dropped; improper handling |
| Burrs | Rough, raised edges on metal surfaces | Impact damage; improper sharpening; wear |
| Chips/nicks | Missing pieces from cutting edges or tips | Impact with hard surfaces; improper use |
| Dull cutting edges | Edges appear rounded or don't cut cleanly | Normal wear; improper use |
| Worn box locks | Excessive looseness in hinge; wobble | Normal wear; overuse |
| Stripped ratchets | Ratchet teeth worn down; won't hold | Normal wear; forcing mechanisms |

## Insulation Damage (Electrosurgical Instruments)

Insulated instruments used in electrosurgery require special inspection for insulation integrity.

- Cracks in coating: May allow current leakage

- Chips or peeling: Exposed metal can cause burns

- Discoloration: May indicate heat damage to insulation

- Bare metal visible: Immediate removal from service

⚠️ Insulation Testing Visual inspection alone is insufficient for insulated instruments. Dedicated insulation testing devices can detect defects not visible to the eye. Follow facility protocol for insulation testing requirements.

# 3.5 Instrument Identification Methods

You will eventually learn to identify thousands of instruments. This takes time and practice. Here's how to identify instruments you don't yet recognize.

## Identification Markings

### Common Marking Locations

- Shank or shaft: Area between handle and working end

- Handle: On one or both handles of hinged instruments

- Blade: On retractor blades or flat surfaces

### Types of Markings

| MARKING | INFORMATION PROVIDED | EXAMPLE |
| --- | --- | --- |
| Manufacturer name/logo | Who made the instrument | Aesculap, V. Mueller, Sklar, Codman |
| Catalog/reference number | Manufacturer's part number | BH123, 34-1234 |
| Instrument name | Common or formal name | Mayo, Metzenbaum, Deaver |
| Size | Length or dimension | 5", 7", 14cm |
| Country of origin | Where manufactured | Germany, USA, Pakistan |
| Facility markings | Internal tracking codes | Colored tape, etching, barcode |

## Identification Resources

### Count Sheets / Pick Lists

The most immediate resource — shows exactly what instruments should be in a specific tray with images or descriptions.

### Instrument Catalogs

Manufacturer catalogs (print or online) contain images and specifications for their instruments. Match markings to find identification.

### Instrument Books/References

Reference books organize instruments by type, specialty, or manufacturer. Good for learning and identification.

### Digital Resources

- Manufacturer websites: Searchable catalogs

- Instrument tracking systems: May include images and specifications

- Specialty apps: Some apps designed for instrument identification

### Ask a Colleague

Experienced technicians are valuable resources. When you don't recognize an instrument, ask. This is how you learn.

🎯 SBD Communication If you cannot identify an instrument, use the standard phrase: "Clarification needed on [description of item]. Please verify identification." Never guess. Ask.

# 3.6 Basic Instrument Categories

Understanding instrument categories helps you identify unfamiliar instruments by recognizing their function. Instruments are generally categorized by what they do.

| CATEGORY | FUNCTION | CHARACTERISTICS | COMMON EXAMPLES |
| --- | --- | --- | --- |
| CUTTING & DISSECTING | Cut tissue, suture, or other materials | Sharp edges; often scissor-like or blade-like | Scissors, scalpels, osteotomes, curettes |
| CLAMPING & OCCLUDING | Compress tissue; control bleeding; occlude vessels | Hinged with ratchets; variety of jaw patterns | Hemostats, clamps (Kelly, Kocher, Mosquito) |
| GRASPING & HOLDING | Hold tissue, needles, or other items | Hinged or spring-loaded; specialized tips | Forceps, needle holders, towel clips, Allis, Babcock |
| RETRACTING & EXPOSING | Hold incision open; expose surgical site | Blade-like or rake-like; hand-held or self-retaining | Deaver, Richardson, Army-Navy, Weitlaner, Balfour |
| PROBING & DILATING | Explore; enlarge openings | Long, slender; various tips | Probes, dilators, sounds |
| SUCTIONING | Remove blood and fluids from surgical site | Hollow tube with tip variations | Yankauer, Poole, Frazier suction tips |
| SUTURING | Hold needles; tie sutures | Specialized jaw surfaces for grip | Needle holders, suture scissors |
| SPECIALTY | Designed for specific procedures | Highly varied | Orthopedic, cardiovascular, neuro, eye instruments |

## Learning Instruments

You will not memorize all instruments immediately. Learning happens through:

- Repetition: The more you handle instruments, the more you recognize them

- Association: Link instrument names to their appearance and function

- Tray familiarity: Learn which instruments go together in common trays

- Active learning: Study instrument books and resources outside work

- Asking questions: Never hesitate to ask about unfamiliar items

💡 Tip: Learn by Category When you encounter an unfamiliar instrument, first ask: "What does this do?" • Does it cut? → Cutting instrument • Does it clamp or hold? → Clamping/grasping instrument • Does it hold an incision open? → Retractor Categorizing helps you understand and remember instruments faster.

# 3.7 Removing Items from Service

Part of your responsibility is to identify instruments that should NOT be used on patients. Removing defective items from service protects patients.

## Criteria for Removal

Remove an instrument from service if it has any of the following:

### Immediate Removal — Do Not Use

- Visible soil that cannot be removed: Indicates possible bioburden

- Cracks: May break during use; harbors bacteria

- Severe corrosion or pitting: Cannot be adequately sterilized

- Broken or missing parts: Instrument incomplete

- Insulation damage (electrosurgical): Risk of patient burns

- Bent or severely misaligned: Won't function properly

- Ratchets won't hold: Clamps won't stay closed

- Sharp burrs or edges that could injure: Safety hazard

### Consider for Repair — May Be Repairable

- Dull scissors: May be sharpened

- Loose box locks: May be tightened

- Minor staining: May be cosmetic

- Worn ratchets: May be refinished

- Minor alignment issues: May be realigned

## Removal Process

- Separate immediately: Remove from clean instruments; do not include in tray

- Tag or label: Use facility-designated method to mark for repair or discard

- Document: Record the item, problem, date, and your initials

- Report: Notify supervisor if unusual or if pattern observed

- Route appropriately: To repair vendor or disposal per facility protocol

🎯 SBD Communication for Defective Items Use standard language when reporting defective instruments: "Issue identified: [instrument name] from [tray/location] — [specific defect]. Removing from service." Example: "Issue identified: Mayo scissors from Major Basic tray — tips misaligned, won't cut cleanly. Removing from service."

## Tracking Removals

Tracking which instruments are removed helps identify:

- Instruments needing frequent repair (may need replacement)

- Problem manufacturers or product lines

- Training issues (improper use patterns)

- Processing issues (damage during reprocessing)

Your facility will have a system for tracking removed instruments. Use it consistently.

# 3.8 Lubrication & Instrument Care

Proper lubrication extends instrument life and maintains function. Lubrication is typically performed after cleaning and before sterilization.

## Why Lubricate

- Reduces friction: Moving parts operate smoothly

- Prevents stiffness: Instruments don't become hard to open/close

- Protects against corrosion: Creates protective barrier

- Extends instrument life: Reduces wear on moving parts

## Instrument Lubricants

Only use lubricants specifically designed for surgical instruments. These are water-soluble, non-toxic, and compatible with the sterilization process.

⚠️ Never Use • Industrial oils or lubricants • Silicone sprays (unless specifically approved) • WD-40 or similar products • Any lubricant not labeled for surgical instrument use Industrial lubricants can leave residue, interfere with sterilization, and harm patients.

## Lubrication Procedure

### Dip Method (Instrument Milk)

- Prepare lubricant per manufacturer instructions (correct dilution, temperature)

- Ensure instruments are clean and free of detergent residue

- Open all hinged instruments

- Immerse instruments completely in lubricant solution

- Work hinges open and closed while submerged

- Remove and allow to drain — do not rinse

- Proceed to sterilization

### Spray Method

- Ensure instrument is clean and dry

- Apply lubricant spray to moving parts (box locks, hinges, ratchets)

- Work mechanism to distribute lubricant

- Wipe excess with lint-free cloth

- Proceed to packaging/sterilization

## Which Instruments to Lubricate

- All hinged instruments (scissors, hemostats, needle holders, clamps)

- Instruments with ratchets

- Instruments with moving parts (retractors with ratchets, etc.)

- Power instrument attachments (per IFU)

### When to Lubricate

- After each decontamination cycle (most effective)

- Per facility protocol (may be every other cycle for some items)

- When instruments feel stiff or difficult to operate

# Knowledge Check: Module 3

Answer these questions to verify your understanding of this module. Discuss any questions with your preceptor.

### Section 3.1 — Purpose of Inspection

- What are the three questions you should ask about every instrument during inspection?

1. _________________________________________________________________________

2. _________________________________________________________________________

3. _________________________________________________________________________

### Section 3.2 — Visual Inspection

- Why is magnification essential for inspection?

_____________________________________________________________________________

_____________________________________________________________________________

- What should you do if you find an instrument with visible soil after decontamination?

_____________________________________________________________________________

### Section 3.3 — Functional Testing

- Describe how to test a scissor for proper function:

_____________________________________________________________________________

_____________________________________________________________________________

- What should you check when testing a ratchet mechanism?

_____________________________________________________________________________

### Section 3.4 — Damage & Defects

- List three types of corrosion damage and their causes:

1. Type: _________________ Cause: _________________________________________

2. Type: _________________ Cause: _________________________________________

3. Type: _________________ Cause: _________________________________________

### Section 3.5 — Identification

- What resources can you use to identify an unfamiliar instrument?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 3.6 — Categories

- Match each instrument category to its function:

Cutting: ___________________________________________________________________

Clamping: _________________________________________________________________

Retracting: ________________________________________________________________

Grasping: _________________________________________________________________

### Section 3.7 — Removing from Service

- List four criteria that require immediate removal of an instrument from service:

1. _________________________________________________________________________

2. _________________________________________________________________________

3. _________________________________________________________________________

4. _________________________________________________________________________

### Section 3.8 — Lubrication

- Why should you never use industrial lubricants on surgical instruments?

_____________________________________________________________________________

_____________________________________________________________________________

# Skills Validation: Module 3

Your preceptor will observe you performing the following skills. Each skill must be demonstrated correctly.

| SKILL | MET | NOT MET | DATE | INITIALS |
| --- | --- | --- | --- | --- |
| Uses proper lighting and magnification for inspection | ☐ | ☐ |  |  |
| Follows systematic inspection sequence | ☐ | ☐ |  |  |
| Correctly identifies clean vs. soiled instruments | ☐ | ☐ |  |  |
| Tests hinged instruments (box lock, ratchet, alignment) | ☐ | ☐ |  |  |
| Performs scissor cutting test correctly | ☐ | ☐ |  |  |
| Identifies common types of instrument damage | ☐ | ☐ |  |  |
| Uses resources to identify unfamiliar instruments | ☐ | ☐ |  |  |
| Correctly categorizes instruments by function | ☐ | ☐ |  |  |
| Properly removes defective items from service | ☐ | ☐ |  |  |
| Applies instrument lubricant correctly | ☐ | ☐ |  |  |
| Documents inspection findings appropriately | ☐ | ☐ |  |  |

### Module Completion

| Learner Name: | ________________________________________ |
| --- | --- |
| Learner Signature: | ________________________________________ |
| Date: | ________________________________________ |
| Preceptor Name: | ________________________________________ |
| Preceptor Signature: | ________________________________________ |
| Date: | ________________________________________ |

# Module 3 Summary

✓ Key Takeaways ✓ Inspection verifies cleanliness, function, and condition ✓ Always use magnification — naked eye inspection is insufficient ✓ Ask three questions: Is it clean? Does it work? Is it safe? ✓ Functional testing confirms instruments work as designed ✓ Learn to identify damage: corrosion, cracks, misalignment, dull edges ✓ Use available resources to identify unfamiliar instruments ✓ Remove defective instruments immediately — document and report ✓ Lubricate with instrument-grade products only

## What's Next

In Module 4, you will learn assembly and tray building — how to correctly assemble instrument sets according to count sheets and prepare them for sterilization. You will learn:

- Using count sheets and pick lists

- Tray organization principles

- Instrument protection during assembly

- Container systems

- Quality checkpoints during assembly

════════════════════════════════════════

Your eyes are the final barrier before an instrument reaches a patient.

Look closely. Question everything. Never settle for "good enough."

════════════════════════════════════════