SPD TECHNICIAN

TRAINING GUIDE

────────────────────────────────────────

MODULE 5

PACKAGING &

WRAPPING

Protecting Sterility from Sterilizer to Patient

IN THIS MODULE • Purpose of Packaging • Types of Sterilization Wrap • Wrapping Techniques • Peel Pouches & Self-Seal Packaging • Chemical Indicators • Labeling Requirements • Package Integrity & Event-Related Sterility • Common Packaging Errors

SIPS HEALTHCARE SOLUTIONS

Sterile By Design™ Curriculum

📚 LEARNING OBJECTIVES

After completing this module, you will be able to: 1. Explain the purpose of packaging in the sterilization process 2. Select appropriate packaging materials for different items and sterilization methods 3. Perform sequential (double) wrapping using proper technique 4. Prepare peel pouches and self-seal packaging correctly 5. Place and interpret chemical indicators appropriately 6. Apply required labeling to all packages 7. Evaluate package integrity before and after sterilization 8. Identify and correct common packaging errors

❤️ THE PATIENT CONNECTION Sterilization kills microorganisms — but packaging keeps them out afterward. A perfectly sterilized instrument becomes contaminated the moment its package is compromised. Your wrapping technique is the last line of defense between the sterilizer and the patient.

# 5.1 Purpose of Packaging

Packaging serves multiple critical functions in the sterilization process. Understanding these functions helps you appreciate why proper technique matters.

## Functions of Packaging

| FUNCTION | HOW IT WORKS |
| --- | --- |
| Allow sterilant penetration | Packaging must allow steam, gas, or other sterilant to reach all surfaces of the contents |
| Maintain sterility after processing | Once sterilized, packaging provides a microbial barrier that keeps contents sterile until use |
| Protect contents | Packaging protects instruments from physical damage, moisture, and environmental contamination during storage and transport |
| Allow aseptic presentation | Packaging must allow contents to be removed without contamination (aseptic transfer to sterile field) |
| Provide identification | Labels on packaging identify contents, sterilization date, and other essential information |

## The Sterile Barrier System

The complete package — wrap, container, or pouch plus any closures and indicators — is called the sterile barrier system (SBS). This system must:

- Withstand the sterilization process without damage

- Remain intact during handling, storage, and transport

- Be opened without contaminating the contents

💡 Two Barriers Are Better Than One Most facilities require double-layer packaging (sequential wrapping) for wrapped items. If the outer layer is compromised, the inner layer still protects sterility. Single-layer packaging may be acceptable for some items — follow your facility's policy.

# 5.2 Types of Sterilization Wrap

Several types of wrapping materials are used in sterile processing. Each has specific properties and appropriate uses.

## Woven (Reusable) Wraps

Traditional fabric wraps that can be laundered and reused.

| TYPE | DESCRIPTION | CONSIDERATIONS |
| --- | --- | --- |
| Muslin (140-thread count) | Traditional cotton fabric; requires 4 layers minimum | Rarely used today; poor barrier properties; labor-intensive |
| Barrier cloth | Tightly woven fabric with better barrier properties than muslin | Must be inspected for holes; limited reuse cycles |

## Nonwoven (Disposable) Wraps

Single-use wraps made from bonded fibers. Most commonly used today.

| TYPE | COMPOSITION | PROPERTIES |
| --- | --- | --- |
| SMS (Spunbond-Meltblown-Spunbond) | Multiple layers of polypropylene fibers | Excellent barrier; good strength; fluid-resistant; most common type |
| Spunbond polypropylene | Single layer of bonded polypropylene fibers | Less expensive; used for outer wrap or light-duty applications |
| Cellulose/synthetic blends | Combination of wood pulp and synthetic fibers | Good conformability; absorbent; varies by manufacturer |

## Wrap Selection Considerations

- Sterilization method: Verify wrap is compatible with your sterilization method (steam, H₂O₂, EO)

- Item size and weight: Heavier items need stronger wraps

- Item shape: Complex shapes may need more conformable material

- Storage conditions: Consider humidity, handling, shelf life requirements

- Manufacturer IFU: Always follow wrap manufacturer's instructions

📌 Nonwoven Is Standard Most facilities use nonwoven (SMS) wraps for routine packaging. They provide consistent barrier properties, require less preparation than woven materials, and eliminate inspection/laundering requirements. Your facility will specify which wraps to use for different applications.

# 5.3 Wrapping Techniques

Proper wrapping technique ensures the package can be sterilized effectively and opened without contamination. Two primary techniques are used.

## Selecting Wrap Size

The wrap must be large enough to completely enclose the item with adequate material for secure folds.

💡 The Wrap Size Rule Wrap should be large enough that the item can be placed diagonally in the center, with each corner able to fold completely over the item and overlap in the center with several inches to spare. Too small = inadequate coverage, compromised barrier Too large = excessive bulk, difficult to handle, may not fit in sterilizer

## Sequential (Double) Wrapping

Two wraps are applied one after the other. Each wrap provides a complete barrier layer.

### Envelope Fold (Parallel Fold)

Most common method for square or rectangular items and trays.

| ENVELOPE FOLD — STEP BY STEP |  |
| --- | --- |
| STEP 1 | Place wrap flat on clean surface, positioned as a diamond (corner pointing toward you) |
| STEP 2 | Place item in center of wrap, positioned squarely |
| STEP 3 | Bring the bottom corner up and over the item; fold back a small cuff (tab) at the top edge |
| STEP 4 | Fold one side corner over the item; tuck under slightly; create a cuff (tab) |
| STEP 5 | Fold the opposite side corner over; tuck under; create a cuff (tab) |
| STEP 6 | Bring the top corner down and over; tuck under the package OR secure with tape |
| STEP 7 | Repeat Steps 1-6 with second wrap for sequential (double) wrapping |

### Square Fold

Alternative method sometimes used for smaller or odd-shaped items.

| SQUARE FOLD — STEP BY STEP |  |
| --- | --- |
| STEP 1 | Place wrap flat, positioned as a square (sides parallel to table edge) |
| STEP 2 | Place item in center |
| STEP 3 | Fold the edge closest to you over the item; fold back a cuff |
| STEP 4 | Fold opposite edge over; fold back a cuff |
| STEP 5 | Fold one side over; tuck ends; fold back cuff |
| STEP 6 | Fold final side over; secure with tape |

## The Purpose of Cuffs (Tabs)

The folded-back cuffs created during wrapping serve a critical purpose:

- Aseptic opening: Allows the package to be opened without touching the sterile inner surface

- Memory: Helps the wrap stay open once unfolded

- Identification: Makes it clear how to open the package properly

⚠️ Cuffs Are Not Optional Every fold must include a cuff (turned-back edge). Without cuffs, the person opening the package must touch potentially contaminated outer surfaces while trying to access sterile contents. Practice until cuffs become automatic.

# 5.4 Peel Pouches & Self-Seal Packaging

Peel pouches (also called peel packs or sterilization pouches) are pre-made packages for individual instruments or small sets.

## Peel Pouch Construction

Standard peel pouches have:

- Paper/film side: Medical-grade paper that allows sterilant penetration

- Plastic side: Clear film for visibility of contents

- Seal area: Either heat-seal or self-seal adhesive strip

- Integrated indicator: Chemical indicator printed on the paper/film (external indicator)

## Selecting Pouch Size

The pouch must be:

- Large enough for the item to fit without forcing

- Small enough that the item doesn't shift excessively

- At least 1 inch larger than the item on all sides

- Long enough to allow proper seal away from the item

💡 The 1-Inch Rule There should be at least 1 inch (2.5 cm) of clearance between the item and all edges of the pouch, including the seal area. This ensures: • Adequate seal integrity • Steam/sterilant can circulate around the item • The item doesn't puncture the pouch during handling

## Peel Pouch Preparation

### Self-Seal Pouches

- Select correct size: Item fits with 1" clearance on all sides

- Place item carefully: Insert without touching the seal area; handle goes in first for hinged instruments

- Position item: Away from seal area; centered in pouch

- Remove adhesive backing: Peel away the protective strip from the adhesive

- Seal the pouch: Press adhesive firmly; run finger along entire seal; no gaps or channels

- Verify seal: Visually confirm seal is complete with no openings

### Heat-Seal Pouches

- Select correct size: Item fits with 1" clearance on all sides

- Place item: Insert carefully; position away from seal area

- Position in sealer: Place open end in heat sealer; ensure straight alignment

- Seal: Activate heat sealer per manufacturer instructions

- Verify seal: Check seal is complete, continuous, without burns or weak spots

## Double Pouching

Some items require double pouching (one pouch inside another) for:

- Additional protection for delicate items

- Sharp items that might puncture a single pouch

- Extended shelf life requirements

### Double Pouching Technique

- Place item in first (inner) pouch and seal

- Place sealed inner pouch in outer pouch — paper side to paper side or plastic side to plastic side

- Seal outer pouch

- Both pouches must have adequate clearance

🎯 Paper to Paper, Plastic to Plastic When double pouching, align the same materials together. Paper to paper OR plastic to plastic — never paper to plastic. This ensures the sterilant can penetrate both layers effectively.

# 5.5 Chemical Indicators

Chemical indicators (CIs) change color when exposed to sterilization conditions. They provide visual evidence that a package has been through a sterilization process.

## Types of Chemical Indicators

Chemical indicators are classified by the parameters they monitor:

| CLASS | NAME | FUNCTION | USE |
| --- | --- | --- | --- |
| Class 1 | Process Indicator | Shows item was exposed to sterilization process | External indicator on every package (tape, pouch print) |
| Class 2 | Specific Test Indicator | Used for specific tests (e.g., Bowie-Dick) | Daily testing of prevacuum sterilizers |
| Class 3 | Single-Parameter Indicator | Reacts to one parameter (e.g., temperature only) | Limited use; less common today |
| Class 4 | Multi-Parameter Indicator | Reacts to 2+ parameters | Common internal indicator type |
| Class 5 | Integrating Indicator | Reacts to all critical parameters over time | High-level internal indicator; correlates with BI |
| Class 6 | Emulating Indicator | Specific to one sterilization cycle type | Cycle-specific verification |

## Using Chemical Indicators

### External Indicators (Class 1)

Placed on the outside of every package to show it has been through the sterilizer.

- Indicator tape: Tape with chemical indicator stripes; used to seal wrapped packages

- Pouch indicators: Printed directly on peel pouch paper

- Container locks: Tamper-evident seals with indicator

Purpose: Distinguish processed from unprocessed items. Does NOT confirm sterilization was successful.

### Internal Indicators (Class 4, 5, or 6)

Placed inside every package to verify sterilant reached the contents.

- One indicator inside each package/container

- Positioned where it will be seen when package is opened

- Must be appropriate for the sterilization method

⚠️ Internal Indicator Required Every package must contain an internal chemical indicator. This verifies that sterilization conditions reached the inside of the package. If an internal indicator is missing or shows inadequate exposure — do not use the contents.

## Interpreting Indicators

| INDICATOR RESULT | MEANING | ACTION |
| --- | --- | --- |
| Complete color change | Indicator was exposed to sterilization parameters | Proceed with use (package may be sterile) |
| Partial or no color change | Inadequate exposure to sterilization parameters | DO NOT USE — quarantine and investigate |
| Indicator damaged/unreadable | Cannot verify exposure | DO NOT USE — reprocess |

Important: Chemical indicators do NOT prove sterility. They only show exposure to sterilization conditions. Biological indicators (BIs) are the only way to verify that sterilization actually killed microorganisms.

# 5.6 Labeling Requirements

Every sterile package must be labeled with essential information for identification, traceability, and safe use.

## Required Label Information

| INFORMATION | PURPOSE | EXAMPLE |
| --- | --- | --- |
| Contents description | Identifies what is in the package | Major Basic Tray, Lap Chole Set |
| Sterilization date | When item was sterilized | 2024-10-15 or 10/15/24 |
| Sterilizer ID | Which sterilizer was used | Steam #2, Sterrad #1 |
| Load/cycle number | Which load the item was in | Load 15, Cycle 42 |
| Technician ID (optional per facility) | Who packaged/processed the item | JD, Tech #4 |
| Lot control number (if applicable) | Tracking for implants and high-risk items | Lot: 12345ABC |

## Label Placement

### Wrapped Packages

- Labels typically placed on indicator tape or applied after wrapping

- Must be visible without opening the package

- Should not cover chemical indicators

- Must adhere securely; won't fall off during handling

### Peel Pouches

- Many pouches have pre-printed areas for labeling

- Write on paper side (not plastic) if hand-labeling

- Use markers approved for sterilization (won't run or transfer)

- Label before sealing when possible

### Containers

- Labels on container lid or designated label area

- May use integrated label holders

- Must be visible and legible

💡 Legibility Matters If a label can't be read, the package shouldn't be used. • Write clearly • Use facility-approved labels and markers • Verify labels are complete before the item leaves your workstation

## Labeling Systems

Facilities use various systems for labeling:

- Hand-written labels: Written on tape or label; requires legible handwriting

- Pre-printed labels: Generated by tracking system; applied to package

- Automated labeling: Labels printed and applied as part of workflow

- Barcode/RFID: Scannable codes for tracking; may supplement written labels

# 5.7 Package Integrity & Event-Related Sterility

A sterile package remains sterile only as long as the package barrier is intact. Understanding what compromises sterility is essential.

## Event-Related Sterility

Event-related sterility means that items remain sterile until something happens (an "event") that compromises the package.

Events that compromise sterility include:

- Tears, holes, or punctures in the packaging

- Wet or moisture-damaged packaging

- Broken seals

- Contamination from handling (touched by unclean hands, dropped on floor)

- Environmental contamination (dust, debris)

- Improper storage conditions

🎯 Time vs. Events Older practices used time-based expiration ("good for 30 days"). Current practice: Items remain sterile based on events, not arbitrary time limits. A properly wrapped and stored item can remain sterile indefinitely — until an event compromises the package. However, many facilities still assign shelf life for inventory management. Follow your facility policy.

## Inspecting Package Integrity

### Before Sterilization

Verify before placing packages in the sterilizer:

- Wrapping is complete with no gaps or openings

- Seals are secure (tape, pouch seals)

- Internal indicator is present

- Label is complete and legible

- Package is appropriate size (not too full, not too loose)

### After Sterilization

Verify after removing from sterilizer:

- External indicators show color change

- Package is dry (wet packs must not be used)

- No visible damage (tears, holes, punctures)

- Seals remain intact

- Labels still legible

### Before Use (at Point of Use)

Final verification before opening:

- Package intact — no tears, holes, or moisture

- Seals intact

- External indicator shows exposure

- Internal indicator (when opened) shows exposure

## Wet Packs

A wet pack is any package that has moisture present on or in the packaging after sterilization.

### Why Wet Packs Occur

- Improper loading (items too close together, no drainage)

- Overloading the sterilizer

- Improper drying cycle

- Package too heavy

- Cool room temperature (condensation)

### Wet Pack Policy

⚠️ Wet Packs Are NOT Sterile Moisture creates a pathway for microorganisms. A wet pack CANNOT be considered sterile. If you identify a wet pack: • Do NOT use • Remove from sterile storage • Reprocess completely (repackage and re-sterilize) • Report per facility policy

# 5.8 Common Packaging Errors

Avoiding these common errors will help ensure your packages maintain sterility effectively.

| ERROR | CONSEQUENCE | PREVENTION |
| --- | --- | --- |
| Wrap too small | Inadequate coverage; exposed areas; tears | Select appropriate size wrap; item should fit diagonally with overlap |
| No cuffs/tabs | Cannot open aseptically; contamination risk | Create cuff on every fold |
| Too tight wrapping | May tear during sterilization or handling; poor steam penetration | Wrap securely but not tight; allow some give |
| Too loose wrapping | Package may open; contents shift and damage wrap | Secure folds; use tape appropriately |
| Incomplete seal (peel pouches) | Sterility compromised; contents not protected | Verify seal across entire width; no channels or gaps |
| Item too close to seal area | Puncture risk; poor seal; sterilant penetration issues | Maintain 1" clearance on all sides |
| Missing internal indicator | Cannot verify sterilant penetration; item unusable | Include internal indicator as standard step; verify before closing |
| Incomplete/illegible labeling | Cannot identify contents; no traceability | Complete label before item leaves workstation; write clearly |
| Using damaged wrap | Compromised barrier; holes or weak spots | Inspect wrap before use; discard any damaged material |
| Excessive tape | Creates barrier to steam penetration; difficult to open | Use minimum tape necessary to secure package |

❤️ The Package Is the Promise Your package promises the OR team that the contents are sterile. A compromised package breaks that promise. Every fold, every seal, every label matters.

# Knowledge Check: Module 5

Answer these questions to verify your understanding of this module. Discuss any questions with your preceptor.

### Section 5.1 — Purpose of Packaging

- List three functions of packaging in the sterilization process:

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

### Section 5.2 — Types of Wrap

- What type of wrap is most commonly used in SPD today?

_____________________________________________________________________________

### Section 5.3 — Wrapping Techniques

- What is the purpose of creating cuffs (tabs) during wrapping?

_____________________________________________________________________________

_____________________________________________________________________________

- What is sequential (double) wrapping?

_____________________________________________________________________________

### Section 5.4 — Peel Pouches

- How much clearance should there be between an item and the edges of a peel pouch?

_____________________________________________________________________________

- When double pouching, how should the pouches be oriented?

_____________________________________________________________________________

### Section 5.5 — Chemical Indicators

- What is the difference between an external (Class 1) indicator and an internal indicator?

_____________________________________________________________________________

_____________________________________________________________________________

- Do chemical indicators prove that an item is sterile? Explain.

_____________________________________________________________________________

_____________________________________________________________________________

### Section 5.6 — Labeling

- List four pieces of information that should be on every sterile package label:

1. _________________________________  2. _________________________________

3. _________________________________  4. _________________________________

### Section 5.7 — Package Integrity

- What is "event-related sterility"?

_____________________________________________________________________________

_____________________________________________________________________________

- What should you do if you identify a wet pack?

_____________________________________________________________________________

### Section 5.8 — Common Errors

- List three common packaging errors and how to prevent them:

1. ___________________________________________________________________________

2. ___________________________________________________________________________

3. ___________________________________________________________________________

# Skills Validation: Module 5

Your preceptor will observe you performing the following skills. Each skill must be demonstrated correctly.

| SKILL | MET | NOT MET | DATE | INITIALS |
| --- | --- | --- | --- | --- |
| Selects appropriate wrap size for item | ☐ | ☐ |  |  |
| Performs envelope fold with proper cuffs on all folds | ☐ | ☐ |  |  |
| Completes sequential (double) wrapping correctly | ☐ | ☐ |  |  |
| Secures package with indicator tape appropriately | ☐ | ☐ |  |  |
| Selects correct peel pouch size (1" clearance) | ☐ | ☐ |  |  |
| Prepares peel pouch with complete seal and proper item placement | ☐ | ☐ |  |  |
| Performs double pouching correctly (when applicable) | ☐ | ☐ |  |  |
| Places internal chemical indicator in package | ☐ | ☐ |  |  |
| Applies complete and legible labeling | ☐ | ☐ |  |  |
| Verifies package integrity before releasing | ☐ | ☐ |  |  |
| Identifies and correctly responds to wet pack scenario | ☐ | ☐ |  |  |
| Interprets chemical indicators correctly | ☐ | ☐ |  |  |

### Module Completion

| Learner Name: | ________________________________________ |
| --- | --- |
| Learner Signature: | ________________________________________ |
| Date: | ________________________________________ |
| Preceptor Name: | ________________________________________ |
| Preceptor Signature: | ________________________________________ |
| Date: | ________________________________________ |

# Module 5 Summary

✓ Key Takeaways ✓ Packaging allows sterilization AND maintains sterility afterward ✓ Sequential (double) wrapping provides two barrier layers ✓ Every fold must have a cuff for aseptic opening ✓ Peel pouches need 1" clearance on all sides ✓ Double pouching: paper to paper OR plastic to plastic ✓ Internal indicator required in every package ✓ Chemical indicators show exposure, not sterility ✓ Complete labeling required for traceability ✓ Wet packs are NOT sterile — must be reprocessed ✓ Package integrity determines sterility (event-related)

## What's Next

In Module 6, you will learn sterilization — the methods and processes that kill microorganisms on processed instruments. You will learn:

- Principles of sterilization

- Steam sterilization (autoclaving)

- Low-temperature sterilization methods

- Loading sterilizers correctly

- Cycle monitoring and documentation

- Sterilizer testing and quality assurance

════════════════════════════════════════

The wrap protects what the sterilizer achieved.

Every fold. Every seal. Every package.

════════════════════════════════════════