SPD TECHNICIAN

TRAINING GUIDE

────────────────────────────────────────

MODULE 8

HIGH-LEVEL DISINFECTION

& SPECIALTY PROCESSING

Processing Heat-Sensitive and Complex Devices

IN THIS MODULE • Understanding High-Level Disinfection • Flexible Endoscope Anatomy • Manual Cleaning of Endoscopes • High-Level Disinfectants • Automated Endoscope Reprocessors (AERs) • Drying, Storage & Transport • Documentation & Quality Assurance • Other Specialty Items

SIPS HEALTHCARE SOLUTIONS

Sterile By Design™ Curriculum

📚 LEARNING OBJECTIVES

After completing this module, you will be able to: 1. Explain the difference between high-level disinfection and sterilization 2. Identify the components of flexible endoscopes and their cleaning challenges 3. Perform manual cleaning of flexible endoscopes according to standards 4. Describe high-level disinfectants and their proper use 5. Operate automated endoscope reprocessors (AERs) safely and effectively 6. Apply proper drying and storage procedures for endoscopes 7. Complete required documentation for HLD processes 8. Process other specialty items according to manufacturer IFUs

❤️ THE PATIENT CONNECTION Endoscope-related infections are among the most serious healthcare-associated infections. Improperly reprocessed endoscopes have caused outbreaks of deadly antibiotic-resistant bacteria. Your meticulous attention to every step of endoscope reprocessing directly prevents patient infections and saves lives.

# 8.1 Understanding High-Level Disinfection

High-level disinfection (HLD) destroys all microorganisms except high numbers of bacterial spores. It is used for semi-critical items that contact mucous membranes or non-intact skin.

## HLD vs. Sterilization

|  | HIGH-LEVEL DISINFECTION | STERILIZATION |
| --- | --- | --- |
| Kills | All microorganisms EXCEPT high numbers of bacterial spores | ALL microorganisms INCLUDING bacterial spores |
| Used for | Semi-critical items (contact mucous membranes) | Critical items (enter sterile tissue or vascular system) |
| Examples | Flexible endoscopes, respiratory equipment, some probes | Surgical instruments, implants |
| Methods | Chemical immersion, AER processing | Steam, H₂O₂ plasma, EO, other sterilization methods |

## When HLD Is Appropriate

HLD is used when:

- The device contacts mucous membranes (GI tract, respiratory tract, etc.)

- The device cannot withstand sterilization methods

- The manufacturer IFU specifies HLD

- The device will not enter sterile tissue

⚠️ HLD Is NOT Sterilization High-level disinfection is appropriate for semi-critical items. If an item will enter sterile tissue or the vascular system, it requires STERILIZATION. Some items previously processed by HLD now have sterilization options — always check the IFU.

## The Critical Role of Cleaning

Cleaning is the most important step in endoscope reprocessing.

- HLD cannot penetrate organic matter

- Bioburden shields microorganisms from disinfectant

- Failed cleaning = failed disinfection, regardless of HLD process

- Studies show manual cleaning removes up to 99.9% of bioburden when done correctly

🎯 Cleaning Comes First You cannot disinfect a dirty scope. The most sophisticated AER cannot compensate for inadequate cleaning. Meticulous manual cleaning is the foundation of safe endoscope reprocessing.

# 8.2 Flexible Endoscope Anatomy

Understanding endoscope design helps you understand why reprocessing is challenging and why each step matters.

## Basic Components

| COMPONENT | DESCRIPTION | CLEANING CHALLENGE |
| --- | --- | --- |
| Control body (handle) | Contains angulation controls, valves, ports | Multiple crevices; valve openings collect debris |
| Insertion tube | Flexible tube inserted into patient | Outer surface contacts tissue; may be damaged |
| Distal tip (bending section) | Contains light guide, lens, channel openings | Most distal point; highest bioburden; delicate |
| Channels (internal) | Air/water, suction/biopsy, auxiliary channels | Long, narrow; cannot see inside; biofilm risk |
| Light guide connector | Connects to light source/processor | Electrical connections; specific cleaning needs |
| Valves | Air/water valve, suction valve | Removable; small parts; collect organic matter |

## Types of Flexible Endoscopes

| TYPE | USE | SPECIAL CONSIDERATIONS |
| --- | --- | --- |
| Gastroscope (EGD) | Upper GI tract (esophagus, stomach, duodenum) | High bioburden; standard channels |
| Colonoscope | Lower GI tract (colon) | Very high bioburden; longer insertion tube |
| Bronchoscope | Airways (trachea, bronchi) | Smaller diameter; narrower channels |
| Duodenoscope (ERCP) | Bile/pancreatic ducts via GI tract | Elevator mechanism — extremely difficult to clean; highest risk |
| Enteroscope | Small intestine | Very long; specialized cleaning requirements |
| Cystoscope (flexible) | Bladder | Smaller; may be sterilizable depending on design |
| ⚠️ Duodenoscopes: Highest Risk Duodenoscopes (ERCP scopes) have an elevator mechanism at the distal tip. This elevator is extremely difficult to clean and has been linked to deadly infection outbreaks. These scopes require meticulous attention to IFU, often including elevator-specific brushes and extended cleaning protocols. Some facilities now sterilize duodenoscopes or use disposable options. |

# 8.3 Manual Cleaning of Endoscopes

Manual cleaning must begin immediately after the procedure (point-of-use pre-cleaning) and continue with thorough cleaning in the reprocessing area.

## Point-of-Use Pre-Cleaning

Pre-cleaning begins at the bedside, immediately after the scope is removed from the patient.

### Steps (at point of use)

- Wipe exterior: Use manufacturer-approved cloth/sponge to wipe insertion tube

- Flush channels: Aspirate enzymatic solution through suction channel

- Flush air/water: Flush air/water channel with water or enzymatic solution

- Transport: Place in closed, leak-proof container for transport to reprocessing

🎯 Why Pre-Clean at Bedside? Blood and body fluids begin drying within MINUTES. Dried organic matter is exponentially harder to remove. Pre-cleaning prevents soil from adhering and forming biofilm. This step is CRITICAL to successful reprocessing.

## Manual Cleaning Procedure

In the reprocessing area, perform thorough manual cleaning before HLD or AER processing.

### PPE Requirements

- Fluid-resistant gown

- Face shield or goggles AND mask

- Heavy-duty gloves (chemical-resistant)

- Consider double gloving

### Leak Testing

Perform leak test BEFORE immersing the scope in liquid.

- Connect leak tester: Attach to scope per manufacturer instructions

- Pressurize: Apply pressure as specified

- Submerge and flex: Look for bubbles while flexing bending section

- Evaluate: If bubbles present — scope has a leak; do not proceed; remove from service

⚠️ Failed Leak Test = Out of Service A scope with a leak CANNOT be safely reprocessed. Liquid entering the internal components will: • Damage the scope • Create an environment for microbial growth • Prevent effective disinfection Remove from service and send for repair.

### Manual Cleaning Steps

| MANUAL CLEANING — STEP BY STEP |  |
| --- | --- |
| STEP 1 | Prepare enzymatic detergent per manufacturer instructions (correct dilution, temperature) |
| STEP 2 | Immerse scope completely in detergent solution |
| STEP 3 | Remove and clean all removable parts (valves, caps) — brush and clean separately |
| STEP 4 | Clean exterior surfaces — wipe entire insertion tube and control body with soft cloth/sponge |
| STEP 5 | Brush ALL channels — use appropriate brushes for each channel; brush must exit distal end; repeat until visually clean |
| STEP 6 | Flush ALL channels — force detergent through all channels with syringe or flush device |
| STEP 7 | Clean elevator mechanism (if present) — use elevator-specific brush; multiple passes |
| STEP 8 | Rinse thoroughly — flush all channels with clean water; remove all detergent |
| STEP 9 | Inspect — visually inspect all accessible areas; if soil remains, repeat cleaning |

## Brushing Technique

- Correct brush size: Use brush specified in scope IFU for each channel

- Full passage: Brush must exit the distal end completely

- Multiple passes: Continue until brush comes out clean

- Clean the brush: Rinse brush between passes; dispose or reprocess after use

- All channels: Every channel must be brushed, including auxiliary channels

⚠️ If the Brush Doesn't Exit — STOP If the brush does not pass completely through the channel: • There may be an obstruction or damage • Do NOT force the brush • Remove scope from service • Notify supervisor • Send for evaluation/repair

# 8.4 High-Level Disinfectants

Several chemical agents are approved for high-level disinfection. Selection depends on device compatibility, facility protocols, and regulatory requirements.

## Common HLD Agents

| AGENT | CONTACT TIME | CONSIDERATIONS |
| --- | --- | --- |
| Glutaraldehyde (2%+) | 20-90 min at 20-25°C (varies by product) | Effective but toxic; requires ventilation; sensitizer; requires thorough rinsing |
| Ortho-phthalaldehyde (OPA) | 12 min at 20°C | Faster than glutaraldehyde; stains skin/tissue gray; no activation required |
| Peracetic acid | Varies (5-15 min typical) | Fast-acting; low toxicity; used in automated systems; corrosive to some materials |
| Hydrogen peroxide (accelerated) | 8-30 min (varies by product) | Less toxic than aldehydes; verify compatibility |

## Using HLD Agents Safely

### Personal Protection

- Always wear appropriate PPE (gloves, face protection, gown)

- Work in well-ventilated area

- Know location of eyewash and safety shower

- Review Safety Data Sheet (SDS) for each chemical

### Solution Monitoring

HLD solutions must be tested to ensure adequate concentration.

- Test frequency: Per manufacturer — typically daily before use and/or each use

- Minimum Effective Concentration (MEC): Solution must be at or above MEC

- Documentation: Record test results; date and initial

- Expiration: Discard solution per manufacturer (reuse life and expiration date)

🎯 Test Before Every Use HLD solutions can become diluted or degraded over time. A solution below MEC will NOT achieve high-level disinfection. Testing is not optional — it is essential to patient safety.

## Manual HLD Process

If manual HLD is performed (rather than AER):

- Verify solution: Test concentration; verify not expired; correct temperature

- Immerse completely: All surfaces and channels must contact disinfectant

- Fill all channels: Force disinfectant through every channel; no air bubbles

- Time exposure: Full contact time per manufacturer instructions

- Rinse thoroughly: Remove all disinfectant residue; use sterile/filtered water for final rinse

- Dry: Purge water from channels; dry exterior

# 8.5 Automated Endoscope Reprocessors (AERs)

AERs automate the HLD process, providing standardized exposure times, temperatures, and rinse cycles.

## Benefits of AERs

- Standardization: Consistent process parameters every cycle

- Documentation: Automatic records of cycle parameters

- Reduced exposure: Less technician contact with chemicals

- Efficiency: Frees staff during cycle; often multiple scopes per cycle

⚠️ AERs Do NOT Replace Manual Cleaning The AER performs disinfection and rinsing — NOT cleaning. Thorough manual cleaning MUST occur before AER processing. An AER cannot remove debris from a scope that was not properly cleaned.

## Using an AER

### Pre-Cycle Checks

- Verify disinfectant level and concentration (test per protocol)

- Check water filters and supply

- Inspect AER for any issues

- Review previous cycle printouts if applicable

### Loading the AER

- Connect channels: Attach all channel connectors per scope and AER IFU

- Verify connections: Ensure all channels are connected; no kinks or leaks

- Position scope: Place as specified by AER manufacturer

- Connect adapters: Use correct adapters for scope model

- Select cycle: Choose appropriate cycle for scope type

- Initiate cycle: Start cycle; document load

### Post-Cycle

- Allow cycle to complete fully

- Review cycle printout for any faults or parameter issues

- Remove scope carefully

- Proceed to drying immediately

## AER Cycle Failures

If the AER indicates a cycle failure or fault:

- Do NOT use the scope: It may not be adequately disinfected

- Document the failure: Record fault code and details

- Investigate: Determine cause (connection issue, chemical level, equipment malfunction)

- Reprocess: Return to manual cleaning and restart HLD process

- Report: Notify supervisor; document actions

# 8.6 Drying, Storage & Transport

Proper drying and storage are critical to preventing recontamination of disinfected endoscopes.

## Why Drying Matters

- Moisture supports bacterial growth: Residual water in channels creates ideal environment for microorganisms

- Biofilm risk: Standing water promotes biofilm formation

- Pseudomonas and other waterborne organisms: Thrive in moist environments

## Drying Procedure

- Purge channels: Force air through all channels to expel water

- Use alcohol flush (if per facility protocol): 70-90% isopropyl or ethyl alcohol through channels facilitates drying

- Final air purge: Blow medical-grade compressed air through all channels

- Wipe exterior: Dry all external surfaces with lint-free cloth

- Remove valves: Store with valves removed to allow air circulation

💡 Alcohol Flush — Check Your Protocol Some facilities use 70% alcohol flush to facilitate drying. This is supported by many guidelines but must be: • Approved by scope manufacturer (check IFU) • Followed by complete air drying • Used consistently per facility protocol

## Storage

### Storage Requirements

- Vertical hanging: Store scopes hanging vertically in a ventilated cabinet

- Valves removed: Store with valves detached to allow air circulation

- Caps off: Leave channel ports open

- Protected: Dedicated cabinet away from contamination sources

- Ventilated: Cabinet should allow air circulation; some have HEPA-filtered positive pressure

- Coiled insertion tube: Avoid tight loops that could damage scope

### Storage Time Limits

Facilities may establish maximum hang time before re-disinfection is required.

- No universal standard — varies by facility (commonly 5-7 days, some up to 14 days)

- Follow your facility's policy

- Document storage time; track when scope was last processed

- If hang time exceeded, scope must be reprocessed before use

## Transport

- Transport in clean, covered container

- Do not coil tightly — prevent damage

- Deliver to procedure area just before use

- Never transport soiled and clean scopes together

# 8.7 Documentation & Quality Assurance

Complete documentation enables traceability and demonstrates compliance with reprocessing standards.

## Required Documentation

| DOCUMENT | CONTENT |
| --- | --- |
| Scope identification | Scope serial number or ID; model; date/time processed |
| Patient linkage | Which scope was used on which patient (enables recall) |
| Pre-cleaning | Documentation that bedside pre-cleaning was performed |
| Leak test | Pass/fail; date/time; technician |
| Manual cleaning | Documentation of cleaning steps completed; technician |
| HLD solution test | MEC test result; date/time; solution lot number |
| AER cycle | Cycle printout; AER ID; cycle number; parameters achieved |
| Technician | Who processed the scope (initials or ID) |

## Quality Assurance Testing

### Routine Testing

- HLD solution concentration: Test per manufacturer (typically daily/each use)

- AER self-tests: Review daily; address any faults

- Water quality: Test rinse water per facility protocol

### Surveillance Cultures

Some facilities perform periodic microbiological surveillance:

- Culture of scopes after reprocessing

- Culture of AER rinse water

- Frequency per facility protocol and risk assessment

- Positive cultures require investigation and remediation

❤️ Traceability Saves Lives If an infection outbreak occurs, you must be able to identify: • Which scope was used on which patient • How that scope was processed • Who processed it • What AER and solutions were used Complete documentation is not bureaucracy — it is patient safety.

# 8.8 Other Specialty Items

Beyond flexible endoscopes, SPD may process other items requiring special handling.

## Heat-Sensitive Devices

Items that cannot tolerate steam sterilization require alternative methods.

| ITEM TYPE | TYPICAL METHOD | NOTES |
| --- | --- | --- |
| Rigid scopes (some) | Low-temp sterilization or HLD | Check IFU — some are steam-compatible |
| Power equipment | Low-temp sterilization (H₂O₂, EO) | Follow manufacturer IFU; remove batteries |
| Cameras/light cords | Low-temp sterilization or HLD | Verify compatibility; protect optics |
| Ultrasound probes | HLD (if semi-critical use) | Follow probe and HLD agent IFU |

## Loaner and Consignment Instruments

Loaner instruments (borrowed from vendors for specific cases) and consignment instruments (vendor-owned, stored at facility) require special handling.

### Loaner Instrument Process

- Receive with documentation: Verify contents against packing list; note condition on arrival

- Inspect: Check for damage, completeness, cleanliness

- Obtain IFUs: Must have manufacturer instructions for processing

- Process completely: Full decontamination, assembly, sterilization cycle

- Document: Record processing like any other instrument

- After use: Decontaminate before return per facility policy

⚠️ Loaners Require Full Processing Never assume a loaner instrument is clean or sterile when it arrives. Process ALL loaners through complete decontamination and sterilization before patient use. This applies even if the vendor says items were "pre-cleaned."

## Implants

Implants (devices permanently placed in the body) require the highest level of care.

- Always sterilize: Implants are critical items; HLD is never acceptable

- Biological indicator: Every implant load requires BI

- Wait for results: Many facilities require negative BI before releasing implants

- Lot tracking: Record implant lot number linked to patient

- Handle with care: Avoid touching implant surfaces; use sterile technique

# Knowledge Check: Module 8

Answer these questions to verify your understanding of this module. Discuss any questions with your preceptor.

### Section 8.1 — Understanding HLD

- What is the difference between high-level disinfection and sterilization?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 8.2 — Endoscope Anatomy

- Why are duodenoscopes considered the highest risk for infection transmission?

_____________________________________________________________________________

_____________________________________________________________________________

### Section 8.3 — Manual Cleaning

- Why is point-of-use pre-cleaning so important?

_____________________________________________________________________________

- What should you do if the brush does not exit the distal end of a channel?

_____________________________________________________________________________

### Section 8.4 — HLD Agents

- Why must HLD solutions be tested before use?

_____________________________________________________________________________

### Section 8.5 — AERs

- What must be done BEFORE placing a scope in the AER?

_____________________________________________________________________________

- What should you do if an AER cycle fails?

_____________________________________________________________________________

### Section 8.6 — Drying & Storage

- Why is proper drying critical for endoscopes?

_____________________________________________________________________________

- How should endoscopes be stored after processing?

_____________________________________________________________________________

### Section 8.7 — Documentation

- List four items that should be documented for each endoscope reprocessing cycle:

1. _________________________________  2. _________________________________

3. _________________________________  4. _________________________________

# Skills Validation: Module 8

Your preceptor will observe you performing the following skills. Each skill must be demonstrated correctly.

| SKILL | MET | NOT MET | DATE | INITIALS |
| --- | --- | --- | --- | --- |
| Dons appropriate PPE for endoscope reprocessing | ☐ | ☐ |  |  |
| Performs leak test correctly | ☐ | ☐ |  |  |
| Performs manual cleaning per IFU (brushing, flushing all channels) | ☐ | ☐ |  |  |
| Tests HLD solution concentration correctly | ☐ | ☐ |  |  |
| Loads AER correctly (all channel connections) | ☐ | ☐ |  |  |
| Reviews AER cycle printout and identifies pass/fail | ☐ | ☐ |  |  |
| Performs proper drying procedure | ☐ | ☐ |  |  |
| Stores endoscope correctly (vertical, valves removed, ventilated) | ☐ | ☐ |  |  |
| Completes all required documentation | ☐ | ☐ |  |  |
| Responds appropriately to failed leak test or AER cycle | ☐ | ☐ |  |  |

### Module Completion

| Learner Name: | ________________________________________ |
| --- | --- |
| Learner Signature: | ________________________________________ |
| Date: | ________________________________________ |
| Preceptor Name: | ________________________________________ |
| Preceptor Signature: | ________________________________________ |
| Date: | ________________________________________ |

# Module 8 Summary

✓ Key Takeaways ✓ HLD kills all microorganisms except high numbers of bacterial spores ✓ HLD is for semi-critical items (contact mucous membranes); sterilization is for critical items ✓ Cleaning is the MOST important step — you cannot disinfect a dirty scope ✓ Pre-cleaning must happen at bedside immediately after use ✓ Leak test BEFORE immersing in liquid — failed leak = out of service ✓ Brush ALL channels; brush must exit distal end ✓ Test HLD solution concentration before EVERY use ✓ AERs automate HLD but do NOT replace manual cleaning ✓ Dry thoroughly — moisture supports bacterial growth ✓ Store vertically, valves removed, in ventilated cabinet ✓ Document everything for traceability

## What's Next

In Module 9, you will learn about quality assurance and process improvement — how to monitor performance, identify problems, and participate in improving SPD operations. You will learn:

- Quality indicators and metrics

- Root cause analysis basics

- Reporting and responding to errors

- Continuous improvement principles

- Your role in quality culture

════════════════════════════════════════

Every step matters. Every channel. Every scope.

Meticulous reprocessing prevents patient infections.

════════════════════════════════════════