SELF-STUDY TRAINING MANUAL

FLEXIBLE ENDOSCOPE

REPROCESSING

A Complete Guide to High-Level Disinfection

Based on AAMI ST91, SGNA Standards & Manufacturer IFUs

This manual is designed for independent study.

Read each chapter in sequence. Complete all knowledge checks.

You will be prepared to safely reprocess flexible endoscopes.

SIPS Healthcare Solutions

Sterile By Design™

HOW TO USE THIS MANUAL

This manual teaches the complete endoscope reprocessing workflow from pre-cleaning through storage. It is designed to be read cover-to-cover without an instructor.

Visual Guide

🛑 STOP - Critical Safety Issue Information that, if ignored, can result in patient harm or equipment damage. Never proceed past a STOP box without understanding and compliance.

⚠️ WARNING - Important Caution Significant information that affects safety or quality. Pay close attention.

💡 KEY CONCEPT Fundamental principle you must understand. These boxes explain the "why" behind procedures.

✓ CHECKPOINT Verification point before proceeding. Confirm all items before moving to the next step.

📋 [FACILITY-SPECIFIC] - Items marked like this require your facility's specific policy, procedure, or value. Ask your supervisor for this information.

Learning Approach

• Read each chapter completely before moving to the next

• Complete the Knowledge Check at the end of each chapter

• If you cannot confidently check every box, re-read the chapter

• After completing all chapters, take the final assessment

• Hands-on competency training with a preceptor is required before independent practice

TABLE OF CONTENTS

Chapter 1: Why Endoscope Reprocessing Matters ................................ 4

Chapter 2: Understanding Endoscope Anatomy .................................. 7

Chapter 3: Personal Protective Equipment .................................... 10

Chapter 4: Pre-Cleaning at Point of Use .................................... 12

Chapter 5: Transport to Reprocessing ........................................ 15

Chapter 6: Leak Testing ..................................................... 17

Chapter 7: Manual Cleaning .................................................. 21

Chapter 8: High-Level Disinfection .......................................... 28

Chapter 9: Rinsing .......................................................... 33

Chapter 10: Drying .......................................................... 35

Chapter 11: Storage ......................................................... 38

Chapter 12: Documentation & Traceability .................................... 41

Chapter 13: Troubleshooting ................................................. 44

Final Assessment ............................................................ 47

Quick Reference: Complete Workflow .......................................... 50

CHAPTER 1
WHY ENDOSCOPE REPROCESSING MATTERS

THE PATIENT SAFETY IMPERATIVE

Flexible endoscopes are inserted into body cavities and come into direct contact with mucous membranes. When contaminated endoscopes are used on patients, serious infections occur-including multi-drug resistant organisms that can be fatal.

Between 2010 and 2023, hundreds of patients were infected by inadequately reprocessed endoscopes. Some died. These were not equipment failures-they were reprocessing failures. The scopes looked clean. They had been through automated reprocessors. But critical steps were missed or performed incorrectly.

🛑 LIVES ARE AT STAKE A single contaminated endoscope can infect dozens of patients before the problem is discovered. Outbreaks have closed endoscopy units, resulted in massive lawsuits, and ended careers. Your attention to every step of reprocessing directly affects whether patients live or die.

WHY ENDOSCOPES ARE DIFFERENT

Most surgical instruments can be steam sterilized in an autoclave. Flexible endoscopes cannot-the heat and pressure would destroy their delicate optics and electronics.

Instead, endoscopes require High-Level Disinfection (HLD), which kills all microorganisms except high numbers of bacterial spores. This level is acceptable because endoscopes contact mucous membranes, which are not sterile.

The challenge is the endoscope design itself:

• Long, narrow channels impossible to see inside

• Complex internal lumens that trap organic material

• Delicate components that limit cleaning methods

• Elevator mechanisms (duodenoscopes) with hidden crevices

💡 THE FUNDAMENTAL PRINCIPLE High-level disinfection only works on CLEAN surfaces. If organic material (blood, tissue, secretions) remains, disinfectant cannot reach the microorganisms underneath. Cleaning is not preparation for disinfection-cleaning IS the foundation of safe reprocessing.

REGULATORY FRAMEWORK

Endoscope reprocessing is governed by multiple standards and guidelines:

• AAMI ST91 - Flexible and semi-rigid endoscope processing in health care facilities

• SGNA - Standards of Infection Prevention in Reprocessing Flexible GI Endoscopes

• CDC - Guideline for Disinfection and Sterilization in Healthcare Facilities

• FDA - Guidance on reprocessing of reusable medical devices

• Manufacturer IFUs - Instructions For Use specific to each scope model

⚠️ IFU IS PRIMARY When manufacturer instructions conflict with general guidelines, follow the manufacturer IFU. The IFU is validated for that specific device. Always have the current IFU available during reprocessing.

YOUR RESPONSIBILITY

You are the last line of defense between a contaminated device and a patient. There are no shortcuts in endoscope reprocessing:

• Every channel must be brushed-every time

• Every step must be completed-no skipping

• Every parameter must be met-no approximations

• Every scope must be documented-no exceptions

When you do this job correctly, patients stay safe. When corners are cut, patients get hurt.

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I understand that reprocessing failures have caused patient deaths

☐ I understand that cleaning must occur before disinfection can work

☐ I understand that the manufacturer IFU is the primary reference

☐ I accept personal responsibility for every scope I reprocess

CHAPTER 2
UNDERSTANDING ENDOSCOPE ANATOMY

WHY ANATOMY MATTERS

You cannot properly clean something you do not understand. Every component of an endoscope requires specific attention during reprocessing. Missing any area creates infection risk.

MAJOR COMPONENTS

Control Handle (Control Body)

The section the operator holds during procedures. Contains:

• Angulation knobs - Control bending section movement

• Air/Water button - Activates insufflation and lens washing

• Suction button - Activates suction through the channel

• Biopsy port - Where instruments enter the working channel

• Electrical connector - Connects to video processor

The buttons contain valves that must be removed and cleaned separately. These valves are a common source of contamination when cleaning is inadequate.

Insertion Tube

The long, flexible tube inserted into the patient. Contains:

• Outer sheath - Smooth surface that contacts patient tissue

• Internal channels - Run the full length of the tube

• Bending section - Near the distal tip, allows angulation

The bending section is articulated and can be damaged by sharp bending or excessive force. Handle with care.

Distal Tip

The working end that enters the patient first. Contains:

• Objective lens - Camera that captures the image

• Light guides - Illuminate the viewing area

• Air/Water nozzle - Insufflates and cleans the lens

• Instrument channel opening - Where biopsy forceps exit

• Elevator (duodenoscopes only) - Raises/lowers accessories

🛑 DUODENOSCOPE ELEVATOR The elevator mechanism on duodenoscopes has caused multiple infection outbreaks. It has crevices that are extremely difficult to clean. If your facility uses duodenoscopes, additional cleaning steps are required. Never assume a duodenoscope elevator is clean.

Umbilical Cord (Light Guide Tube)

Connects the endoscope to the light source and processor:

• Contains electrical cables for video signal

• Contains channels that connect to processor

• Connector must be protected during reprocessing

INTERNAL CHANNELS

Channels are where contamination hides. You cannot see inside them. Every channel must be cleaned.

| Channel | Purpose | Cleaning Priority |
| --- | --- | --- |
| Air Channel | Insufflates body cavity | High - Narrow lumen |
| Water Channel | Flushes lens, irrigates | High - Narrow lumen |
| Suction/Biopsy Channel | Suction and instrument passage | Critical - Largest exposure to patient material |
| Auxiliary Water (if present) | Additional irrigation | High - Follow IFU |
| Elevator Channel (duodenoscopes) | Elevator mechanism access | Critical - Most difficult to clean |

💡 CHANNEL CONFIGURATION VARIES Different endoscope models have different channel configurations. Some have separate air and water channels; some have a combined air/water channel. Some have auxiliary channels. Always consult the specific IFU to identify all channels.

📋 [FACILITY-SPECIFIC] Obtain the IFU for each endoscope model your department uses. Document the channel configuration for each.

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I can identify the control handle, insertion tube, distal tip, and umbilical cord

☐ I understand that all internal channels must be cleaned

☐ I understand the duodenoscope elevator requires special attention

☐ I know to consult the IFU for scope-specific channel configurations

CHAPTER 3
PERSONAL PROTECTIVE EQUIPMENT

PROTECTING YOURSELF

Endoscope reprocessing exposes you to bloodborne pathogens, chemical hazards, and infectious material. Proper PPE is required throughout the reprocessing workflow.

REQUIRED PPE

| PPE Item | Requirement | Purpose |
| --- | --- | --- |
| Gloves | Chemical-resistant, appropriate for disinfectant used | Protects from chemicals and biological material |
| Gown | Fluid-resistant, long-sleeved, tied at back | Protects clothing and skin from splashes |
| Eye Protection | Face shield OR goggles with side shields | Protects eyes from splashes and aerosols |
| Mask | Surgical mask minimum; N95 if aerosol risk | Protects respiratory system |
| Shoe Covers | Fluid-resistant | Protects footwear, prevents tracking contamination |

WHEN TO WEAR WHAT

Full PPE Required

• Pre-cleaning at point of use

• Transporting contaminated scopes

• Leak testing

• Manual cleaning

• Loading/unloading AER

• Handling high-level disinfectants

Gloves and Gown Required

• Handling scopes after disinfection

• Rinsing

• Drying

• Placing scopes in storage

⚠️ CHEMICAL HAZARDS High-level disinfectants (glutaraldehyde, OPA, peracetic acid) are hazardous. They can cause respiratory sensitization, chemical burns, and eye damage. Ensure adequate ventilation. Know the location of eyewash stations and spill kits.

DONNING AND DOFFING

Donning Order (Putting On)

• 1. Hand hygiene

• 2. Gown - Tie at neck and waist

• 3. Mask - Secure over nose and mouth

• 4. Eye protection - Ensure seal around eyes

• 5. Gloves - Pull cuffs over gown sleeves

Doffing Order (Removing)

• 1. Gloves - Peel off without touching exterior

• 2. Eye protection - Remove by touching only earpieces/headband

• 3. Gown - Unfasten ties, roll away from body, dispose

• 4. Mask - Remove by ear loops only, never touch front

• 5. Hand hygiene - Immediately

💡 GLOVE CHANGES Change gloves: between contaminated and clean activities, when torn or punctured, when moving from one scope to another, after contact with high-level disinfectants.

📋 [FACILITY-SPECIFIC] Know the location of PPE supplies, eyewash stations, and chemical spill kits in your reprocessing area.

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I know what PPE is required for each reprocessing step

☐ I know the correct order for donning and doffing PPE

☐ I understand the chemical hazards of high-level disinfectants

☐ I know when to change gloves

CHAPTER 4
PRE-CLEANING AT POINT OF USE

WHY PRE-CLEANING CANNOT WAIT

Pre-cleaning begins IMMEDIATELY after the endoscope is removed from the patient-while still in the procedure room. This is the most time-critical step in reprocessing.

Organic material (blood, mucus, tissue, feces) begins drying within minutes. Once dried, this material becomes extremely difficult to remove. Biofilm can begin forming in as little as 20 minutes.

🛑 TIME IS CRITICAL Pre-cleaning must begin within minutes of procedure completion. There is no acceptable delay. A scope that sits uncleaned becomes a scope that may not be safely reprocessable. Every minute matters.

PRE-CLEANING PROCEDURE

STEP 1: WIPE THE INSERTION TUBE

While the scope is still connected to the light source, use a soft lint-free cloth moistened with enzymatic detergent. Wipe from the control handle toward the distal tip. Remove all visible debris.

STEP 2: FLUSH SUCTION CHANNEL

Place the distal tip in enzymatic detergent solution. Activate suction to draw solution through the suction channel. Continue until solution runs clear.

📋 [FACILITY-SPECIFIC] Minimum suction volume: _____ mL per manufacturer IFU

STEP 3: FLUSH AIR/WATER CHANNELS

Depress air and water buttons to flush detergent solution through these channels. Alternate between air and water.

STEP 4: DISCONNECT FROM PROCESSOR

Remove the scope from the light source and video processor.

STEP 5: ATTACH PROTECTIVE CAP

Cover the electrical connector with the water-resistant cap to prevent damage during transport and cleaning.

STEP 6: PLACE IN TRANSPORT CONTAINER

Place the scope in an approved, closed transport container.

ENZYMATIC DETERGENT

Enzymatic detergents contain proteins that break down organic material. They are more effective than water alone.

• Use detergent approved for endoscope use

• Prepare fresh solution per manufacturer instructions

• Never reuse enzymatic solution between scopes

• Observe correct dilution and temperature

📋 [FACILITY-SPECIFIC] Approved enzymatic detergent: _____________

📋 [FACILITY-SPECIFIC] Dilution ratio: _____ per manufacturer instructions

💡 PRE-CLEANING IS NOT DISINFECTION Pre-cleaning removes gross contamination. It does not disinfect the scope. A pre-cleaned scope is still contaminated and must be transported and handled as biohazardous.

✓ CHECKPOINT BEFORE TRANSPORT Insertion tube wiped? ☐ Suction channel flushed? ☐ Air/water channels flushed? ☐ Protective cap attached? ☐ Scope in closed container? ☐

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I understand that pre-cleaning must begin immediately after the procedure

☐ I know to wipe the insertion tube from handle toward tip

☐ I know to flush all channels with enzymatic detergent

☐ I understand that pre-cleaned scopes are still contaminated

CHAPTER 5
TRANSPORT TO REPROCESSING

SAFE TRANSPORT

After pre-cleaning, the scope must be transported to the reprocessing area. During transport, you must prevent environmental contamination and protect the scope from damage.

TRANSPORT REQUIREMENTS

• Use a closed, leak-proof container

• Container must be labeled as biohazardous

• Container must be large enough to prevent coiling

• Scope should not contact sides of container

• Transport immediately-do not allow scope to dry

🛑 NEVER TRANSPORT UNCONTAINED Never carry a contaminated endoscope through hallways without a closed container. This spreads pathogens, violates infection control standards, and risks damage to the scope.

TIMING

The scope should arrive at the reprocessing area and begin leak testing as quickly as possible:

📋 [FACILITY-SPECIFIC] Maximum transport time: _____ minutes

If any delay is unavoidable:

• Keep the scope moist with enzymatic-soaked towels

• Document the delay and reason

• Notify reprocessing staff of the delay

💡 MOISTURE IS CRITICAL A moist scope can be cleaned. A dry scope may not be. If pre-cleaned material dries in the channels, it may bond to surfaces and become impossible to remove. Preventing drying is essential.

HANDOFF COMMUNICATION

When delivering the scope to reprocessing, communicate:

• Scope identifier (serial number or facility ID)

• Procedure performed

• Any problems during procedure (scope dropped, difficult cleaning)

• Time pre-cleaning was completed

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I know to use a closed, labeled container for transport

☐ I understand why the scope must not dry during transport

☐ I know what information to communicate during handoff

CHAPTER 6
LEAK TESTING

PURPOSE OF LEAK TESTING

The endoscope's outer sheath protects internal components (optics, electronics) from fluid. If this sheath is punctured, fluid enters the scope during cleaning and disinfection, causing:

• Damage to internal electronics (costly repair)

• Damage to optics (image quality loss)

• Creation of areas where microorganisms can hide

• Progressive damage with each subsequent use

Leak testing detects damage BEFORE immersion. A scope that fails leak testing should never be submerged.

🛑 TEST BEFORE IMMERSION Leak testing is performed BEFORE the scope enters the cleaning sink. If you discover a leak after immersion, you have already caused additional internal damage. The leak test is your protection - and the scope's.

LEAK TESTING PROCEDURE

STEP 1: ATTACH LEAK TESTER

Connect the leak tester to the scope's venting connector. This is a specific port designed for leak testing-consult the IFU for location.

STEP 2: PRESSURIZE

Inflate the scope to the pressure specified in the manufacturer IFU.

📋 [FACILITY-SPECIFIC] Leak test pressure: _____ mmHg per scope IFU

STEP 3: SUBMERGE

While maintaining pressure, completely submerge the scope in water.

STEP 4: OBSERVE FOR BUBBLES

Watch for a continuous stream of bubbles from any location. Check:
• Entire insertion tube
• Bending section (flex in all directions while observing)
• Distal tip
• Control body and buttons
• All seams and connections

STEP 5: MAINTAIN OBSERVATION

Continue observing for the time specified in the IFU.

📋 [FACILITY-SPECIFIC] Observation time: _____ seconds per scope IFU

STEP 6: RECORD RESULT

Document PASS or FAIL.

INTERPRETING RESULTS

| Observation | Interpretation | Action |
| --- | --- | --- |
| No bubbles, pressure holds | PASS | Proceed to cleaning |
| Steady stream of bubbles from any location | FAIL | Remove from service immediately |
| Single bubble, then stops | Possible trapped air | Retest-if uncertain, treat as FAIL |
| Pressure drops but no visible bubbles | Possible slow leak | Check connections, retest |
| Bubbles from channel opening | NORMAL | Channels are not sealed-this is expected |

IF LEAK TEST FAILS

🛑 STOP - DO NOT CONTINUE If the scope fails the leak test: 1. Remove from water IMMEDIATELY 2. Do NOT continue cleaning or disinfection 3. Dry the exterior to prevent further water intrusion 4. Tag as "Out of Service - Leak Test Failure" 5. Notify supervisor 6. Initiate repair process per facility protocol

A scope that fails leak testing cannot be safely reprocessed. There are no exceptions. Attempting to clean or disinfect a leaking scope causes additional damage and creates infection risk.

COMMON LEAK LOCATIONS

• Bending section - Most common; flexed repeatedly during procedures

• Distal tip - Contact with instruments and tissue

• Insertion tube - Bite damage, crushing in doors/drawers

• Control body buttons - Repeated use causes wear

✓ CHECKPOINT BEFORE PROCEEDING Leak tester properly attached? ☐ Correct pressure achieved? ☐ Scope fully submerged? ☐ Bending section flexed during observation? ☐ Full observation time completed? ☐ Result documented? ☐

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I understand why leak testing must occur before immersion

☐ I know the proper sequence for leak testing

☐ I can interpret leak test results

☐ I know exactly what to do if a scope fails the leak test

CHAPTER 7
MANUAL CLEANING

THE MOST CRITICAL STEP

Manual cleaning is the foundation of safe reprocessing. No amount of high-level disinfection can compensate for inadequate cleaning. Studies consistently show that 15-30% of endoscopes fail visual cleanliness checks-meaning cleaning is often done poorly.

🛑 CLEANING IS NON-NEGOTIABLE If organic material remains after cleaning, high-level disinfectant cannot reach the surface beneath. Microorganisms survive under the bioburden, and patients become infected. There is no shortcut. There is no substitute. Every channel must be cleaned.

EQUIPMENT AND SUPPLIES

• Clean sink dedicated to endoscope cleaning

• Fresh enzymatic detergent at correct dilution

• Channel cleaning brushes (correct sizes per IFU)

• Cleaning adapters for channel flushing

• Syringes or powered irrigation device

• Lint-free cloths or sponges

• Clean rinse water

• Full PPE

⚠️ SINGLE-USE BRUSHES Many facilities use single-use brushes. Check your facility policy. If reusable brushes are used, they must be cleaned, inspected for damage, and high-level disinfected or sterilized between uses per manufacturer instructions.

📋 [FACILITY-SPECIFIC] Brush type: ☐ Single-use  ☐ Reusable with HLD/sterilization

PREPARE THE CLEANING SOLUTION

• Mix enzymatic detergent fresh for each scope

• Follow manufacturer dilution instructions exactly

• Use water at the temperature specified by detergent manufacturer

• Never reuse enzymatic solution-it becomes saturated and ineffective

📋 [FACILITY-SPECIFIC] Enzymatic detergent: _____________

📋 [FACILITY-SPECIFIC] Dilution: _____________

📋 [FACILITY-SPECIFIC] Water temperature: _____________

CLEANING PROCEDURE

STEP 1: IMMERSE AND SOAK

Submerge the scope completely in enzymatic detergent solution. Ensure all channels fill with solution (no air pockets). Soak for the time specified by detergent manufacturer.

📋 [FACILITY-SPECIFIC] Minimum soak time: _____ minutes

STEP 2: CLEAN EXTERNAL SURFACES

Using a soft cloth or sponge saturated with detergent:
• Wipe the entire insertion tube from handle to tip
• Clean the control body, including around buttons
• Clean the distal tip thoroughly
• Clean the umbilical cord
• Pay attention to all crevices and seams

STEP 3: REMOVE AND CLEAN VALVES

Remove all detachable valves and buttons:
• Disassemble per manufacturer IFU
• Clean all surfaces with brush and detergent
• Brush internal ports and openings
• Set aside in detergent solution

STEP 4: BRUSH ALL CHANNELS

This is the most critical cleaning step.

💡 BRUSHING TECHNIQUE For each channel: 1. Select the correct brush size (must contact channel walls) 2. Insert brush at one end 3. Advance until brush exits the other end 4. Withdraw brush while rotating 5. Inspect brush for debris 6. Clean brush in detergent 7. Repeat until brush emerges clean (minimum 3 passes) 8. Never force a brush-if resistance is met, stop and investigate

Every channel must be brushed. There are no exceptions. Channels that "look clean" still harbor invisible contamination.

📋 [FACILITY-SPECIFIC] Document brush sizes for each scope model in your department.

STEP 5: FLUSH ALL CHANNELS

Using syringes or powered irrigation:
• Flush each channel with enzymatic detergent
• Use the volume and pressure specified in the IFU
• Ensure solution exits all channel openings

📋 [FACILITY-SPECIFIC] Minimum flush volume per channel: _____ mL

STEP 6: CLEAN ELEVATOR (DUODENOSCOPES ONLY)

If processing a duodenoscope:
• Raise and lower elevator repeatedly while submerged
• Brush around and under elevator with specific elevator brush
• Flush elevator channel with detergent
• This step is critical-elevator contamination has caused outbreaks

STEP 7: RINSE THOROUGHLY

Rinse all channels and surfaces with clean water to remove detergent and loosened debris.

VISUAL INSPECTION

After cleaning, visually inspect the scope:

• Examine the distal tip lens-should be clear

• Check all channel openings for visible debris

• Inspect the bending section for smooth movement

• Look for any damage (cuts, cracks, discoloration)

🛑 VISIBLE DEBRIS = RECLEAN If you see any visible debris after cleaning, the scope is not adequately cleaned. Return to brushing and flushing. Do NOT proceed to high-level disinfection.

COMMON CLEANING ERRORS

• Insufficient brushing - biofilm remains in channels

• Wrong brush size - brush doesn't contact channel walls

• Not brushing all channels - assuming some are "clean"

• Reusing enzymatic solution - saturated solution is ineffective

• Insufficient flush volume - debris remains in channels

• Not cleaning valves - major contamination source

• Rushing - inadequate contact time and attention

✓ CHECKPOINT BEFORE HLD Enzymatic soak completed? ☐ All external surfaces cleaned? ☐ All valves removed and cleaned? ☐ Every channel brushed until clean? ☐ All channels flushed? ☐ Elevator cleaned (if duodenoscope)? ☐ Visual inspection passed? ☐

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I understand that cleaning is the most critical reprocessing step

☐ I know to brush every channel until the brush emerges clean

☐ I understand why enzymatic solution must not be reused

☐ I know to stop and reclean if visible debris remains

CHAPTER 8
HIGH-LEVEL DISINFECTION

WHAT HLD ACHIEVES

High-Level Disinfection (HLD) kills all microorganisms except high numbers of bacterial spores. This level of disinfection is required for semi-critical devices-those that contact mucous membranes or non-intact skin.

HLD is not sterilization. Some bacterial spores may survive. This is acceptable for endoscopes because mucous membranes are not sterile.

💡 HLD REQUIREMENTS For HLD to be effective, four conditions must be met: 1. The scope must be thoroughly CLEANED first 2. Disinfectant must be at or above Minimum Effective Concentration (MEC) 3. Full contact time must be achieved 4. Correct temperature must be maintained

HLD METHODS

Automated Endoscope Reprocessor (AER)

Most facilities use AERs for HLD. Benefits:

• Standardized, reproducible process

• Automatic channel perfusion

• Temperature and time control

• Cycle documentation

• Reduced staff chemical exposure

AER does NOT clean the scope. Manual cleaning must be completed BEFORE loading.

Manual HLD

When AER is unavailable, manual HLD may be performed. This requires meticulous technique and documentation.

COMMON HIGH-LEVEL DISINFECTANTS

| Disinfectant | Typical Contact Time | Key Considerations |
| --- | --- | --- |
| Glutaraldehyde (≥2%) | 20-90 min at 20-25°C | Respiratory sensitizer; requires ventilation |
| OPA (Ortho-phthalaldehyde) | 10-12 min at 20°C | Stains proteins gray; monitor for anaphylaxis |
| Peracetic acid | Per manufacturer | Fast acting; check material compatibility |
| Hydrogen peroxide solutions | Per manufacturer | Various formulations; follow specific IFU |

⚠️ CHEMICAL COMPATIBILITY Not all disinfectants are compatible with all endoscopes. Verify compatibility in BOTH the disinfectant IFU AND the endoscope IFU before use.

📋 [FACILITY-SPECIFIC] Approved HLD chemical: _____________

📋 [FACILITY-SPECIFIC] Contact time: _____________

📋 [FACILITY-SPECIFIC] Temperature requirement: _____________

MINIMUM EFFECTIVE CONCENTRATION (MEC)

Disinfectant concentration decreases with use and time. Before EVERY use, you must verify the solution meets the Minimum Effective Concentration (MEC).

• Test solution with the manufacturer's test strip

• Compare to the MEC threshold on the test strip bottle

• If below MEC, discard the solution and prepare fresh

• Document the test result

🛑 NEVER USE SUB-MEC SOLUTION Solution below MEC will not achieve high-level disinfection. Microorganisms survive. Patients become infected. There are no exceptions-if MEC fails, replace the solution.

AER PROCEDURE

STEP 1: VERIFY AER STATUS

Confirm the AER has passed its self-test and MEC testing is current.

STEP 2: LOAD THE SCOPE

Place the scope in the AER basin:
• Connect all channel adapters to appropriate ports
• Ensure no kinks in the insertion tube
• Verify all connections are secure

⚠️ ALL CHANNELS CONNECTED If a channel is not connected to the AER, disinfectant does not flow through it. That channel remains contaminated. Verify every connection before starting the cycle.

STEP 3: LOAD VALVES AND COMPONENTS

Place valves and other detachable components in the AER per manufacturer instructions.

STEP 4: CLOSE AND START CYCLE

Close the AER and initiate the appropriate cycle.

STEP 5: MONITOR CYCLE

Remain available during the cycle:
• Do not open the AER during the cycle
• Respond to any alarms immediately
• Do not assume completion-verify cycle completed successfully

STEP 6: VERIFY COMPLETION

Confirm the cycle completed without errors. Print or document the cycle record.

MANUAL HLD PROCEDURE

If AER is unavailable:

STEP 1: TEST CONCENTRATION

Verify disinfectant meets MEC using test strip.

STEP 2: IMMERSE COMPLETELY

Submerge the scope with no air pockets.

STEP 3: FILL ALL CHANNELS

Use syringes to fill every channel with disinfectant. Confirm solution exits all openings.

STEP 4: MAINTAIN CONTACT TIME

Set a timer. Maintain immersion for the FULL contact time specified. Do not remove early for any reason.

STEP 5: MAINTAIN TEMPERATURE

Monitor temperature if required by the disinfectant.

STEP 6: DOCUMENT

Record: chemical used, lot number, MEC test result, contact time, temperature.

🛑 CONTACT TIME IS MINIMUM The specified contact time is the MINIMUM required to kill microorganisms. Removing the scope even one minute early can leave viable pathogens. Set a timer. Do not estimate.

✓ CHECKPOINT AFTER HLD MEC verified before cycle? ☐ All channels connected/filled? ☐ Full contact time achieved? ☐ Temperature maintained (if required)? ☐ Cycle completed without errors? ☐ Result documented? ☐

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I understand that cleaning must occur before HLD

☐ I know to test disinfectant concentration before every use

☐ I understand that all channels must contact the disinfectant

☐ I know that contact time is a minimum, not a target

CHAPTER 9
RINSING

PURPOSE OF RINSING

After high-level disinfection, all disinfectant residue must be removed from the scope. Residual disinfectant can cause:

• Mucosal irritation or burns to the next patient

• Chemical damage to scope components

• Interference with subsequent procedures

RINSE WATER QUALITY

The quality of rinse water matters. Tap water may contain microorganisms that recontaminate the scope.

| Rinse Stage | Acceptable Water Quality |
| --- | --- |
| First rinse (remove bulk disinfectant) | Tap water may be acceptable per IFU |
| Final rinse | Per IFU-may require sterile, filtered, or bacteria-free water |

⚠️ FOLLOW THE IFU If the disinfectant or endoscope IFU specifies sterile water for final rinse, using tap water creates recontamination risk. Follow the IFU exactly.

📋 [FACILITY-SPECIFIC] Final rinse water specification: _____________

📋 [FACILITY-SPECIFIC] Source of rinse water: _____________

RINSE PROCEDURE

STEP 1: FLUSH ALL CHANNELS

Flush each channel with appropriate rinse water. Use the volume specified in the IFU.

STEP 2: RINSE EXTERIOR SURFACES

Rinse all external surfaces to remove disinfectant.

STEP 3: REPEAT IF REQUIRED

Some disinfectants require multiple rinse cycles. Follow the disinfectant IFU.

📋 [FACILITY-SPECIFIC] Number of rinse cycles required: _____________

📋 [FACILITY-SPECIFIC] Rinse volume per channel: _____________

STEP 4: RINSE VALVES AND COMPONENTS

Rinse all detachable parts that were disinfected.

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I understand why rinsing is necessary after HLD

☐ I know the rinse water quality requirements for my facility

☐ I know to follow the IFU for number of rinse cycles

CHAPTER 10
DRYING

WHY DRYING IS CRITICAL

Moisture enables microbial growth. A properly disinfected endoscope that is stored wet will become contaminated. Drying is not optional-it is essential to maintaining the disinfected state.

🛑 WET STORAGE = CONTAMINATION Microorganisms need moisture to survive and multiply. Even a perfectly disinfected scope will grow bacteria if stored wet. Within hours, waterborne organisms can multiply to dangerous levels. Dry completely before storage.

DRYING PROCEDURE

Alcohol Flush

Flushing channels with alcohol displaces water and aids evaporation:

STEP 1: PREPARE ALCOHOL

Use 70% isopropyl alcohol or 70% ethyl alcohol.

STEP 2: FLUSH ALL CHANNELS

Pass alcohol through every channel using syringes or irrigation adapters.

📋 [FACILITY-SPECIFIC] Alcohol flush volume per channel: _____ mL

STEP 3: PURGE WITH AIR

Follow alcohol flush with forced air to remove the alcohol.

⚠️ ALCOHOL COMPATIBILITY Verify alcohol flush is compatible with your endoscope model. Some manufacturers have specific requirements. Check the IFU.

Forced Air Drying

STEP 1: CONNECT TO AIR SOURCE

Use filtered, compressed air or connect to a drying cabinet.

STEP 2: DRY ALL CHANNELS

Force air through every channel for the time specified in your facility protocol.

📋 [FACILITY-SPECIFIC] Minimum forced air drying time: _____ minutes

STEP 3: DRY EXTERIOR

Wipe exterior surfaces with a lint-free cloth.

STEP 4: VERIFY DRYING

Visually confirm no moisture is visible. Channels should expel only air, no droplets.

Drying Cabinets

Many facilities use dedicated drying cabinets that provide:

• HEPA-filtered air circulation

• Controlled environment

• Direct channel connections

• Documentation of drying time

If your facility uses drying cabinets, follow the manufacturer's instructions for connection and minimum drying time.

📋 [FACILITY-SPECIFIC] Drying cabinet model: _____________

📋 [FACILITY-SPECIFIC] Minimum cabinet drying time: _____________

✓ CHECKPOINT BEFORE STORAGE Alcohol flush completed? ☐ All channels purged with air? ☐ Minimum drying time achieved? ☐ No visible moisture on exterior? ☐ No droplets from channels? ☐

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I understand that wet storage causes recontamination

☐ I know the drying procedure for my facility

☐ I know to verify channels are completely dry before storage

CHAPTER 11
STORAGE

STORAGE REQUIREMENTS

How you store endoscopes directly affects whether they remain safe for patient use. Improper storage can undo all your reprocessing work.

Vertical Hanging

• Store scopes hanging vertically

• Distal tip points down to allow drainage

• Control handle at top

• Insertion tube hangs freely-do not coil

Ventilated Cabinet

• Dedicated cabinet for endoscope storage only

• HEPA-filtered air circulation preferred

• Clean, dry, dust-free environment

• Individual positions for each scope-no contact between scopes

Valves Removed

• Store valves separately (not attached to scope)

• This allows air circulation through channels

• Prevents moisture accumulation

🛑 NEVER STORE SCOPES: • In transport containers • In closed cases or trays • Coiled in drawers • Touching other scopes • With valves attached • Wet or damp These practices cause recontamination.

REPROCESSING WINDOW (HANG TIME)

A reprocessed endoscope does not stay disinfected indefinitely. Your facility defines a maximum storage time-the "reprocessing window" or "hang time"-after which the scope must be reprocessed again before use.

| Organization | Recommendation |
| --- | --- |
| SGNA | Reprocess if stored >7 days |
| Many facilities | 24-72 hours |
| Your facility | [Per facility policy] |

📋 [FACILITY-SPECIFIC] Maximum storage time before reprocessing: _____________

⚠️ EXPIRED = REPROCESS If a scope has exceeded the reprocessing window, it must be reprocessed before patient use-even if it was properly processed initially. Do not use an expired scope.

STORAGE LABELING

Each stored scope should be labeled with:

• Scope identifier (serial number or facility code)

• Date and time reprocessing completed

• Expiration date/time (when reprocessing window ends)

• Initials of person who reprocessed

CABINET MAINTENANCE

• Clean cabinet interior regularly per facility protocol

• Replace HEPA filters per manufacturer schedule

• Document cleaning and maintenance

📋 [FACILITY-SPECIFIC] Storage cabinet cleaning frequency: _____________

✓ CHECKPOINT - PROPER STORAGE Scope hanging vertically? ☐ Distal tip down? ☐ Not coiled? ☐ Not touching other scopes? ☐ Valves removed? ☐ Completely dry? ☐ Labeled with expiration? ☐ In ventilated cabinet? ☐

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I know to store scopes hanging vertically with valves removed

☐ I understand why coiled storage causes contamination

☐ I know my facility's reprocessing window

☐ I know to reprocess if the storage window is exceeded

CHAPTER 12
DOCUMENTATION & TRACEABILITY

WHY DOCUMENTATION MATTERS

Documentation proves you did the work correctly. When a patient develops an infection or a surveyor asks questions, your documentation is your evidence.

💡 IF IT'S NOT DOCUMENTED, IT DIDN'T HAPPEN In healthcare, documentation is proof. Without records, you cannot demonstrate compliance, investigate problems, or defend your practice. Document everything.

WHAT TO DOCUMENT

For every reprocessing cycle, record:

| Step | Document |
| --- | --- |
| Pre-cleaning | Time completed, initials |
| Leak test | Pass/Fail, scope identifier |
| Manual cleaning | Completion verified, initials |
| HLD | Chemical, lot #, MEC test result, cycle #, pass/fail |
| Drying | Method, completion time |
| Storage | Date, time, location, expiration |
| Patient use | Patient ID, procedure, date, scope identifier |

TRACEABILITY

You must be able to trace in both directions:

Scope → Patients

Given a scope identifier, you can find every patient on whom it was used.

Patient → Scope

Given a patient, you can find which scope was used and its complete reprocessing history.

💡 INVESTIGATION SCENARIO A patient develops an infection 3 days after endoscopy. Investigators need to know: • Which scope was used • Who reprocessed it and when • Were all steps completed correctly • What other patients were exposed to that scope Your documentation answers these questions.

DOCUMENTATION QUALITY

• Complete - Every required field filled

• Accurate - Times and values are correct

• Legible - Can be read by anyone

• Timely - Recorded at time of activity, not later

• Indelible - Cannot be erased or altered without detection

If you make an error, draw a single line through it, write the correction, initial, and date. Never erase or white-out.

RECORD RETENTION

📋 [FACILITY-SPECIFIC] Reprocessing record retention period: _____ years

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I understand that documentation is proof of compliance

☐ I know what must be documented for each reprocessing cycle

☐ I understand scope-to-patient traceability

☐ I know how to correct documentation errors properly

CHAPTER 13
TROUBLESHOOTING

COMMON PROBLEMS

Leak Test Failure

| Situation | Action |
| --- | --- |
| Bubbles from insertion tube or bending section | FAIL - Remove from service, dry exterior, tag out, notify supervisor, initiate repair |
| Bubbles from channel opening | NORMAL - Channels are not sealed |
| Pressure won't hold, no visible bubbles | Check all connections, retest; if still fails, treat as failure |
| Uncertain result | Retest; if still uncertain, treat as failure |

Cleaning Issues

| Situation | Action |
| --- | --- |
| Visible debris after cleaning | Reclean - Return to brushing step |
| Brush won't pass through channel | Check brush size, check for blockage, do not force; if blocked, remove from service |
| ATP fails (if used) | Reclean - Manual cleaning was inadequate |
| Scope was not pre-cleaned | Document deviation, extend enzymatic soak, clean thoroughly |

HLD Issues

| Situation | Action |
| --- | --- |
| MEC test fails | Discard solution, prepare fresh, retest before use |
| AER cycle aborts/errors | Review alarm code, troubleshoot per AER manual, reprocess from beginning |
| Contact time interrupted | Restart entire HLD cycle |
| Wrong chemical used | Stop, consult supervisor, may need to remove scope from service |

Storage Issues

| Situation | Action |
| --- | --- |
| Scope exceeds reprocessing window | Must reprocess before patient use |
| Scope found stored wet | Reprocess - assume contaminated |
| Scope found coiled | Inspect for damage, reprocess |
| Scope stored with valves attached | Reprocess - channels may have moisture |

WHEN TO ESCALATE

Notify your supervisor immediately for:

• Repeated leak test failures on the same scope

• Suspected infection linked to endoscopy

• AER malfunctions

• Chemical spills or exposure

• Any situation where you are unsure of the correct action

💡 WHEN IN DOUBT, ASK If you are ever uncertain whether a scope is safe for patient use, STOP and ask your supervisor. No one will criticize you for being cautious. The consequences of using a contaminated scope are severe.

KNOWLEDGE CHECK - Before proceeding, confirm you understand:

☐ I know what to do if a scope fails the leak test

☐ I know what to do if cleaning appears inadequate

☐ I know what to do if a scope exceeds the reprocessing window

☐ I know when to escalate to my supervisor

FINAL ASSESSMENT

Answer all questions. You should be able to answer every question correctly before performing endoscope reprocessing independently.

Section A: True or False

1. Pre-cleaning can be delayed until the scope reaches the reprocessing area. _____

2. Leak testing is performed after manual cleaning. _____

3. Manual cleaning is the most critical step in reprocessing. _____

4. If a scope looks clean, brushing is optional. _____

5. HLD contact time can be shortened if the scope was cleaned thoroughly. _____

6. A scope stored wet may become contaminated. _____

7. Scopes should be stored hanging vertically with valves removed. _____

8. If a scope exceeds the reprocessing window, it can still be used in an emergency. _____

Section B: Fill in the Blank

9. Pre-cleaning must begin _____________ after procedure completion.

10. If a scope fails the leak test, it must be _____________ from service.

11. Every channel must be brushed until the brush comes out _____________.

12. MEC stands for _____________ _____________ _____________.

13. Alcohol flush before storage helps _____________ the channels.

14. Scopes should be stored with the _____________ tip pointing down.

Section C: Short Answer

15. List four things that must be documented for each reprocessing cycle.

_________________________________________________________________

_________________________________________________________________

16. What do you do if you see visible debris on a scope after completing manual cleaning?

_________________________________________________________________

_________________________________________________________________

17. Why must endoscopes be stored dry?

_________________________________________________________________

_________________________________________________________________

18. What are three situations that require immediate supervisor notification?

_________________________________________________________________

_________________________________________________________________

ANSWER KEY

Section A

1. FALSE - Pre-cleaning must begin immediately

2. FALSE - Leak testing is performed BEFORE cleaning (before immersion)

3. TRUE

4. FALSE - You cannot see inside channels; always brush

5. FALSE - Contact time is a minimum; never shorten

6. TRUE

7. TRUE

8. FALSE - Expired scopes must always be reprocessed first

Section B

9. immediately

10. removed

11. clean

12. Minimum Effective Concentration

13. dry

14. distal

Section C

15. Any four: Scope identifier, date/time, person performing, leak test result, cleaning verification, HLD parameters (chemical, lot #, MEC result, cycle/time), drying, storage location/expiration, patient use

16. Return to manual cleaning. Reclean and re-brush until clean. Do NOT proceed to HLD.

17. Moisture enables microbial growth. A wet scope will become contaminated even after proper HLD.

18. Any three: Repeated leak test failures, suspected infection linked to endoscopy, AER malfunctions, chemical spills/exposure, uncertain situations

QUICK REFERENCE: COMPLETE WORKFLOW

STEP 1: PRE-CLEAN

Immediately at point of use. Wipe insertion tube, flush channels with enzymatic detergent, attach protective cap, place in transport container.

STEP 2: TRANSPORT

Closed, labeled container. Deliver promptly-do not allow scope to dry.

STEP 3: LEAK TEST

Before immersion. Pressurize, submerge, observe for bubbles. PASS → proceed. FAIL → remove from service.

STEP 4: MANUAL CLEAN

Fresh enzymatic solution. Clean exterior, remove/clean valves, brush EVERY channel until clean, flush all channels, visual inspection.

STEP 5: HIGH-LEVEL DISINFECT

Verify MEC. Connect all channels (AER) or fill all channels (manual). Complete full contact time. Document.

STEP 6: RINSE

Remove all disinfectant. Use appropriate water quality per IFU. Rinse all channels and surfaces.

STEP 7: DRY

Alcohol flush all channels. Forced air drying. Verify no moisture remains.

STEP 8: STORE

Hang vertically, distal tip down. Valves removed. Ventilated cabinet. Label with expiration.

STEP 9: DOCUMENT

Complete record of all steps, times, results, personnel.

STEP 10: BEFORE PATIENT USE

Verify scope is within reprocessing window. Visual inspection. Attach valves.

Every step. Every channel. Every time.

Patient safety depends on your diligence.

TRAINING COMPLETION

I have completed the Endoscope Reprocessing Self-Study Manual. I have read all chapters, completed all knowledge checks, and passed the final assessment.

I understand that hands-on competency training with a qualified preceptor is required before I can perform endoscope reprocessing independently.

| Learner Printed Name | Date |
| --- | --- |
| Learner Signature | Assessment Score /18 |

END OF SELF-STUDY MANUAL